## FINAL 1&2nd PAGE
Protect Sheet Password : bstech
CERTIFICATE NO. |  |  |  |  | : | BT 140067 | [6 empty] | PAGE: |  | 1 | OF | 3
SALES ORDER NO. |  |  |  |  | : | BSO-14-000511
CUSTOMER |  |  |  |  | : | UNIWELD PRODUCTS (USA) PTE LTD
40 MACTAGGART ROAD
SINGAPORE 368085
INSTRUMENT |  |  |  |  | : | DIGITAL THERMOMETER WITH SENSOR
MANUFACTURER |  |  |  |  | : | FLUKE
MODEL NO. |  |  |  |  | : | 51 II
SERIAL NO. |  |  |  |  | : | 26170097WS
TAG NO. |  |  |  |  | : | -
RANGE |  |  |  |  | : | (-200 to 1372)°C
DATE CALIBRATED |  |  |  |  | : | 2014-01-30
RECOMMENDED DUE DATE |  |  |  |  | : | 2015-01-30
AMBIENT TEMPERATURE |  |  |  |  | : | (21 ± 2) °C
RELATIVE HUMIDITY |  |  |  |  | : | (55 ± 10) % rh
BS Tech´s organisation and practices have been duly accredited and are in compliant to the requirements of ISO 9001.
The reference measurement standards used are traceable to National Metrology Centre,(NMC,SINGAPORE) and/or other National Standards.
MEASUREMENT TRACEABILITY
| REFERENCE STANDARDS | [5 empty] | TAG NO. |  |  | SERIAL NO. |  | CAL DATE |  | DUE DATE
| MULTIFUNCTION CALIBRATOR | [5 empty] | BS 1394 |  |  | 9760013 |  | 2013-08-02 |  | 2014-08-02
| REFERENCE THERMOCOUPLE | [5 empty] | BS 1363 |  |  | - |  | 2013-10-24 |  | 2014-10-24
| PRT SENSOR(PT 100)-4 WIRE | [5 empty] | BS 1340 |  |  | EP 53022 |  | 2013-06-20 |  | 2014-06-20
| PRECISION DMM | [5 empty] | BS 1326 |  |  | US36075261 |  | 2013-10-03 |  | 2014-10-03
CALIBRATED BY : | [8 empty] | APPROVED BY :
|  |  |  | R SENTHIL | [6 empty] | S KARUPPIAH
CERTIFICATE NO. |  |  | : | BT 140067 | [8 empty] | PAGE: |  | 2 | OF | 3
.
METHOD OF CALIBRATION
The calibration certificate was produced in for the points specified in the relevant in-House Technical Procedure,
BST 01-H & BST 05,as defined for the instrument.
The calibration was performed in a controlled temperature source.
The described instrument has been calibrated at | [7 empty] | BS Laboratory under the ambient conditions stated in the first page.
As Received Condition :
Initial testing found the instrument to be in specification for the parameters tested.
As Returned Condition :
At the completion of the service, measured values were in specification for the parameters tested.
RESULTS OF CALIBRATION
♦ | The expanded uncertainty of measurement associated with the calibration is | [9 empty] | estimated at a confidence
level of approximately 95% with a coverage factor of k = 2.
REMARKS
♦ | The results of the calibration are given on the attached following pages.
♦ | The results and their associated uncertainties are applicable at the time of calibration.
♦ | Highlighted values fall out of the manufacturer's specifications. Values not highlighted fall within the range allowed.
♦ | No adjustment was done unless otherwise stated. | [16 empty] | Must, Please to see
♦ | The user should determine the suitability of the instrument for its intended use.
♦ | The reports shall not be reproduced except in full, unless the management representative of BS TECH PTE LTD has given approval in writing and comply with the requirements specified in ISO 9001.
Please to see
Please to see
|  | No adjustment was done unless otherwise stated. | [13 empty] | SERIAL NO.
|  | Adjustment was done, Please refer the data sheets. | [13 empty] | TAG NO.
Initial testing found the instrument to be in specification for the parameters tested.
Initial testing found the instrument to be out of specification for the parameters tested.
♦ This equipment has been calibrated and conforms to the requirements of the manufacturers specification.
♦ The results reported herein have been performed in accordance with the laboratory's terms of accreditation under the Singapore Accreditation Council-Singapore Laboratory Accreditation Scheme.
CUSTOMER
VESSEL NAME
USER NAME
RANGE
CAL RANGE
SET VALUE
TEST VALUE

## cert 2 page
[10 empty] | % rh. | °C | °F
CERTIFICATE NO : | BT 140067 |  |  |  | PAGE: | 3 | OF | 3
TEST RESULTS:
WITH SENSOR
ACTUAL READING |  | INSTRUMENT READING |  | CORRECTION | EXPANDED UNCERTAINTY
|  | BEF ADJ | AFT ADJ
°C
0.14 |  | 0.1 | - | 0.04 | 0.38888
50.35 |  | 50.1 | - | 0.25 | 0.38888
100.48 |  | 100.1 | - | 0.38 | 0.41556
0.00 |  | 0.0 | - | 0.00 | #DIV/0!
0.00 |  | 0.0 | - | 0.00 | #DIV/0!
0.00 |  | 0.0 | - | 0.00 | #DIV/0!
0.00 |  | 0.0 | - | 0.00 | #DIV/0!
0.00 |  | 0.0 | - | 0.00 | #DIV/0! |  |  |  | SETTING TEMPERATURE
79.61 |  | 80.4 | - | -0.79 | 0
99.23 |  | 100.1 | - | -0.87 | 0 |  |  |  | °C
20
25
30
RELATIVE HUMIDITY :
ACTUAL READING |  | INSTRUMENT READING |  | CORRECTION | EXPANDED UNCERTAINTY
|  | BEF ADJ | AFT ADJ
%rh
35.2 |  | 35 | - | 0.2 | 3.5764
55.8 |  | 55 | - | 0.8 | 3.8276
86.4 |  | 85 | - | 1.4 | 3.9728
T2:
ACTUAL READING |  | INSTRUMENT READING |  | CORRECTION | EXPANDED UNCERTAINTY
|  | BEF ADJ | AFT ADJ
°C
0.09 |  | -0.4 | - | 0.49 | 0.38888
50.09 |  | 49.8 | - | 0.29 | 0.38888
100.12 |  | 99.9 | - | 0.22 | 0.41556
150.11 |  | 149.8 | - | 0.31 | #DIV/0!
200.04 |  | 199.7 | - | 0.34 | #DIV/0!
250.23 |  | 249.9 | - | 0.33 | #DIV/0!
K TYPE(WITHOUT SENSOR)
ACTUAL READING |  | INSTRUMENT READING |  | CORRECTION | EXPANDED UNCERTAINTY
|  | BEF ADJ | AFT ADJ
°C
-100.00 |  | -100.2 | - | 0.20 | 0.45662
0.00 |  | -0.2 | - | 0.20 | 0.45662
500.00 |  | 499.8 | - | 0.20 | 0.5143
1000 |  | 999.9 | - | 0.10 | 0.62266
1200 |  | 1200 | - | 0.00 | 0.62266
****** End Of Test Data ******

## worksheet
[5 empty] | Laboratory Worksheet | [18 empty] | RES °C | 1
Certificate no. |  |  |  | : | BT 140067 |  |  |  |  | 12 | [13 empty] | RES % R.H. | FALSE
Certificate no. |  |  |  | : | BT 140067 | [7 empty] | Ambient Temperature |  |  | : | (21 ± 2) °C
Sales Order no. |  |  |  | : | BSO-14-000511 | [7 empty] | Relative Humidity |  |  | : | (55 ± 10) % rh | [5 empty] | UNIT | MODEL | MANUFACTURE | INSTRUMENT | DATE | PAGE NO | DUE DATE | TEMPE
User |  |  |  | : | Uniweld Products (USA) Pte Ltd | [7 empty] | Page |  |  | : | 1 |  | OF | 3 |  |  | °C | SDC-10 | YOKOGAWA | TEMPERATURE CONTROLLER | 2011-01-01 | 3 | 2014-04-30 | (20 ± 1) °C
[5 empty] | 40 Mactaggart Road | [7 empty] | Date Calibrated |  |  | : | 2014-01-30 | [5 empty] | °F | GCT-630 | GW | GROUND CONTINUITY TESTER | 2011-01-02 | 4 | 2014-07-30 | (20 ± 2) °C
[5 empty] | Singapore 368085 | [7 empty] | Recommended Due Date |  |  | : | 2015-01-30 | [5 empty] | %RH | MT 160C | FENWAL | PRESSURE GAUGE | 2011-01-03 | 5 | 2015-01-30 | (20 ± 3) °C
[23 empty] | rpm | HA-250K | WIKA | DIGITAL THERMOMETER | 2011-01-04 | 6 | 2016-01-30 | (21 ± 2) °C
[23 empty] | RTD | DLM-1000 | EIRELEC | MERCURY IN-GLASS THERMOMETER | 2011-01-05 | 7 | - | (21 ± 3) °C
Instrument |  |  |  | : | DIGITAL THERMOMETER WITH SENSOR | [17 empty] | mV | J-221 | ANRITSU | RADIOMETER C/W SENSOR'S | 2011-01-06 | 8 |  | (21 ± 4) °C
Manufacturer |  |  |  | : | FLUKE | [7 empty] | Range |  |  | : | (-200 to 1372)°C |  |  |  |  | Sec | V | 1015 | GOULD BASS | INFRARED THERMOMETER | 2011-01-07 | 9 |  | (22 ± 2) °C
Model no. |  |  |  | : | 51 II | [7 empty] | Tag no. |  |  | : | - | [5 empty] | µA | MP 54 | BROOKSTONE | TEMPERATURE GAGE & PROBE | 2011-01-08 | 10 |  | (22 ± 3) °C
Serial no. |  |  |  | : | 26170097WS | [7 empty] | Part no. |  |  | : | - | [5 empty] | mA | MP 55 | BLAK-RAY | ULTRAVIOLET MEASURING METER | 2011-01-09 | 11 |  | (23 ± 2) °C
[23 empty] | A | KP-15 | JEIC | SOUND LEVEL METER | 2011-01-10 | 12 |  | (23 ± 3) °C
[13 empty] | RESOLUTION °C |  |  | : | 0.1 | [5 empty] | kΩ | 51 II | DOVE | RTD SENSOR | 2011-01-11 | 13 |  | (24 ± 2) °C
[13 empty] | RESOLUTION %r.h. |  |  | : | [6 empty] | MΩ | FT-500 | CAPRI | LOW PRESSURE GAUGE | 2011-01-12 | 14 |  | (24 ± 3) °C
[8 empty] | NON SINGLAS | [14 empty] | Sec | 422 | DANFOSS | HIGH PRESSURE GAUGE | 2011-01-13 | 15 |  | (25 ± 2) °C
[23 empty] | PSI |  | FLUKE | HI-LOW DUAL PRESSURE GAUGE | 2011-01-14 | 16 |  | (25 ± 3) °C
[23 empty] | BAR |  | ONO SOKKI | OIL DIFFERENTIAL PRESSURE SWITCH | 2011-01-15 | 17 |  | (25 ± 4) °C
Reference Standards: | [24 empty] | GOSSEN | THERMOMETER | 2011-01-16 | 18 |  | (26 ± 2) °C
| 1 |  | BS 1394 |  |  | Multifunction Calibrator |  |  |  |  | 9760013 |  | 2013-08-02 |  | 2014-08-02 |  |  | BST 05 |  |  | NOT OK |  |  |  | LINE SEIKI | ADVANCED TACHOMETER | 2011-01-17 |  |  | (27 ± 2) °C
| 2 |  | BS 1363 |  |  | Reference thermocouple |  |  |  |  | - |  | 2013-10-24 |  | 2014-10-24 |  |  | BST 05 |  |  | OK |  |  |  | ALTEK | TEMPERATURE INDICATOR | 2011-01-18 |  |  | (28 ± 2) °C
| 3 |  | BS 1340 |  |  | PRT SENSOR(PT 100)-4 WIRE |  |  |  |  | EP 53022 |  | 2013-06-20 |  | 2014-06-20 |  |  | BST 01-H/L |  |  | NOT OK |  |  |  |  | DIGITAL THERMOMETER | 2011-01-19 |  |  | (29 ± 2) °C
| 4 |  | BS 1326 |  |  | Precision DMM |  |  |  |  | US36075261 |  | 2013-10-03 |  | 2014-10-03 |  |  | BSE 02 or BST 05 |  |  | OK |  |  |  |  | THERMOCOUPLE THERMOMETER (K-type) | 2011-01-20 |  |  | (30 ± 2) °C
| 5 |  |  |  |  | #N/A |  |  |  |  | #N/A |  | #N/A |  | #N/A |  |  | #N/A |  |  | #N/A | [5 empty] | 2011-01-21 |  |  | (31 ± 2) °C
| 6 |  |  |  |  | #N/A |  |  |  |  | #N/A |  | #N/A |  | #N/A |  |  | #N/A |  |  | #N/A | [5 empty] | 2011-01-22 |  |  | (32 ± 2) °C
| PROCEDURE : | [7 empty] | BST 01-H & BST 05 | [17 empty] | 2011-01-23 |  |  | (33 ± 2) °C
| UNCERTAINTY : | [25 empty] | 2011-01-24 |  |  | (34 ± 2) °C
| PLACE OF CALIBRATION : | [7 empty] | BS Laboratory | [17 empty] | 2011-01-25 |  |  | (35 ± 2) °C
2011-01-26
2011-01-27
2011-01-28
2011-01-29
Calibrated by : |  |  |  | R SENTHIL | [9 empty] | Approved By: |  | S KARUPPIAH | [10 empty] | 2011-01-30
2011-01-31
2011-02-01
Doc No : BS - WP ISSUE Date : 23.09.05 Rev Date : 24.06.08/1 | [26 empty] | 2011-02-02
2011-02-03
2011-02-04
2011-02-05
2011-02-06
2011-02-07
2011-02-08
2011-02-09
2011-02-10
2011-02-11
2011-02-12
|  | N SHARGUNA RAJ | [24 empty] | 2011-02-13
|  | T VIJAYBABU | [9 empty] | SINGLAS | [14 empty] | 2011-02-14
|  | R RAMESH | [9 empty] | NON SINGLAS | [14 empty] | 2011-02-15
|  | R SENTHIL |  |  |  |  | % of reading | [19 empty] | 2011-02-16
|  | DIVO S.K |  |  |  |  | % of full scale | [19 empty] | 2011-02-17
|  | R G SEKAR | [24 empty] | 2011-02-18
|  | MAHEN | [24 empty] | 2011-02-19
|  | S SURESH |  |  |  |  | Uniweld Products (USA) Pte Ltd | [19 empty] | 2011-02-20
|  | RAJESH |  |  |  |  | BS Laboratory |  |  |  |  | MULTIRANGE | [14 empty] | 2011-02-21
|  | C BALAMURUGAN | [9 empty] | - | [14 empty] | 2011-02-22
2011-02-23
2011-02-24
|  | M S SANDHU | [24 empty] | 2011-02-25
|  | S KARUPPIAH | [24 empty] | 2011-02-26
|  | REDHY | [24 empty] | 2011-02-27
|  | DIVO S.K | [24 empty] | 2011-02-28
|  | GOBI | [24 empty] | 2011-03-01
|  | T VIJAYBABU | [24 empty] | 2011-03-02
2011-03-03
Section no. : |  |  |  |  | : | [21 empty] | 2011-03-04
2011-03-05
2011-03-06
[7 empty] | Log no. | [19 empty] | 2011-03-07
[7 empty] | ID no. | [19 empty] | 2011-03-08
[7 empty] | Control no. | [19 empty] | 2011-03-09
[7 empty] | Assent no. | [19 empty] | 2011-03-10
[7 empty] | PO no. | [19 empty] | 2011-03-11
[7 empty] | Geaso no. | [19 empty] | 2011-03-12
[7 empty] | Code no. | [19 empty] | 2011-03-13
[7 empty] | Chart no. | [19 empty] | 2011-03-14
[7 empty] | Equipment no. | [19 empty] | 2011-03-15
[7 empty] | Part no. | [19 empty] | 2011-03-16
[7 empty] | Type no. | [19 empty] | 2011-03-17
[7 empty] | Tool no. | [19 empty] | 2011-03-18
[7 empty] | Tag no. | [19 empty] | 2011-03-19
[7 empty] | Location | [19 empty] | 2011-03-20
2011-03-21
2011-03-22
2011-03-23
2011-03-24
2011-03-25
2011-03-26
2011-03-27
2011-03-28
2011-03-29
2011-03-30
2011-03-31
2011-04-01
2011-04-02
2011-04-03
2011-04-04
2011-04-05
2010-04-15
2010-04-16
2010-04-17
2010-04-18
2010-04-19
2010-04-20
2010-04-21
2010-04-22
2010-04-23
2010-04-24
2010-04-25
2010-04-26
2010-04-27
2010-04-28
2010-04-29
2010-04-30
2010-05-01
2010-05-02
2010-05-03
2010-05-04
2010-05-05
2010-05-06
2010-05-07
2010-05-08
2010-05-09
2010-05-10
2010-05-11
2010-05-12
2010-05-13
2010-05-14
2010-05-15
2010-05-16
2010-05-17
2010-05-18
2010-05-19
2010-05-20
2010-05-21
2010-05-22
2010-05-23
2010-05-24
2010-05-25
2010-05-26
2010-05-27
2010-05-28
2010-05-29
2010-05-30
2010-05-31
2010-06-01
2010-06-02
2010-06-03
2010-06-04
2010-06-05
2010-06-06
2010-06-07
2010-06-08
2010-06-09
2010-06-10
2010-06-11
2010-06-12
2010-06-13
2010-06-14
2010-06-15
2010-06-16
2010-06-17
2010-06-18
2010-06-19
2010-06-20
2010-06-21
2010-06-22
2010-06-23
2010-06-24
2010-06-25
2010-06-26
2010-06-27
2010-06-28
2010-06-29
2010-06-30
2010-07-01
2010-07-02
2010-07-03
2010-07-04
2010-07-05
2010-07-06
2010-07-07
2010-07-08
2010-07-09
2010-07-10
2010-07-11
2010-07-12
2010-07-13
2010-07-14
2010-07-15
2010-07-16
2010-07-17
2010-07-18
2010-07-19
2010-07-20
2010-07-21
2010-07-22
2010-07-23
2010-07-24
2010-07-25
2010-07-26
2010-07-27
2010-07-28
Note on F4: TYPE CERT NO ONLY
Note on F5: DO N'T TYPE THIS ROW
Note on D23: Please Type Master Tag No Only Ex: BS 1320
Note on D24: Please Type Master Tag No Only Ex: BS 1319
Note on D25: Please Type Master Tag No Only Ex: BS 1317
Note on D26: Please Type Master Tag No Only Ex: BS 1331
Note on D27: Please Type Master Tag No Only Ex: BS 1391
Note on D28: Please Type Master Tag No Only Ex: BS 1364

## Sensor Raw Data
|  | RAW DATA | [19 empty] | Ambient Temperature |  |  |  |  | : | (21 ± 2) °C
Certificate no. | BT 140067
RAW DATA | [28 empty] | RAW DATA
TEMPERATURE | [10 empty] | DESCRIPTION | SENSOR | [7 empty] | degree | Coverage factor | [7 empty] | OUT TEMPERATURE | [17 empty] | degree | Coverage factor
UNIT °C | [19 empty] | of freedom | k factor | [7 empty] | UNIT °C | [17 empty] | of freedom | k factor
REF 1 | PRT | PRT | PRT | [7 empty] | RTD SENSOR | SPRT
REF 2 | DMRTD | DMRTD | DMRTD | [7 empty] | RTD SENSOR | PRT
REF 3 | LOWBATH | LOWBATH | MIDBATH | [7 empty] | S TYPE TC | STC
REF 4 | [9 empty] | THERMOHYGROMETER(TEMP) |  | THMTEMP
SET POINT °C | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |  | THERMOHYGROMETER(HUM) |  | THMHUM | [16 empty] | SET POINT °C | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
Ref.inst.rdg 1 | 0.14 | 50.35 | 100.48 | [16 empty] | 1 | 13.97 | [7 empty] | Ref.inst.rdg 1 | 0.09 | 50.09 | 100.12 | 150.11 | 200.04 | 250.23 | [11 empty] | 1 | 13.97
Ref.inst.rdg 2 | 0.14 | 50.35 | 100.48 | [8 empty] | INDICATOR | [7 empty] | 2 | 4.53 | [7 empty] | Ref.inst.rdg 2 | 0.09 | 50.09 | 100.12 | 150.11 | 200.04 | 250.23 | [11 empty] | 2 | 4.53
Ref.inst.rdg 3 | 0.14 | 50.35 | 100.48 | [16 empty] | 3 | 3.31 | [7 empty] | Ref.inst.rdg 3 | 0.09 | 50.09 | 100.12 | 150.11 | 200.04 | 250.23 | [11 empty] | 3 | 3.31
Ref.inst.rdg 4 | 0.14 | 50.35 | 100.48 | [7 empty] | DMM-RTD | DMRTD | [7 empty] | 4 | 2.87 | [7 empty] | Ref.inst.rdg 4 | 0.09 | 50.09 | 100.12 | 150.11 | 200.04 | 250.23 | [11 empty] | 4 | 2.87
Ref.inst.rdg 5 | 0.14 | 50.35 | 100.48 | [7 empty] | DMM 81/2(RTD) | DMRTD8.5 | [7 empty] | 5 | 2.65 | [7 empty] | Ref.inst.rdg 5 | 0.09 | 50.09 | 100.12 | 150.11 | 200.04 | 250.23 | [11 empty] | 5 | 2.65
Ref.inst.rdg 6 | 0.14 | 50.35 | 100.48 | [7 empty] | DMM-TC | DMTC | [7 empty] | 6 | 2.52 | [7 empty] | Ref.inst.rdg 6 | 0.09 | 50.09 | 100.12 | 150.11 | 200.04 | 250.23 | [11 empty] | 6 | 2.52
Ref.inst.rdg 7 | 0.14 | 50.35 | 100.48 | [16 empty] | 7 | 2.43 | [7 empty] | Ref.inst.rdg 7 | 0.09 | 50.09 | 100.12 | 150.11 | 200.04 | 250.23 | [11 empty] | 7 | 2.43
Ref.inst.rdg 8 | 0.14 | 50.35 | 100.48 | [7 empty] | TEMP BATH | HEAT SOURCE | [7 empty] | 8 | 2.37 | [7 empty] | Ref.inst.rdg 8 | 0.09 | 50.09 | 100.12 | 150.11 | 200.04 | 250.23 | [11 empty] | 8 | 2.37
Ref.inst.rdg 9 | 0.14 | 50.35 | 100.48 | [16 empty] | 9 | 2.32 | [7 empty] | Ref.inst.rdg 9 | 0.09 | 50.09 | 100.12 | 150.11 | 200.04 | 250.23 | [11 empty] | 9 | 2.32
Ref.inst.rdg 10 | 0.14 | 50.35 | 100.48 | [7 empty] | USING INFRARED ONLY | IR | [7 empty] | 10 | 2.28 | [7 empty] | Ref.inst.rdg 10 | 0.09 | 50.09 | 100.12 | 150.11 | 200.04 | 250.23 | [11 empty] | 10 | 2.28
MEAN TEMP | 0.14 | 50.35 | 100.48 | 0.00 | 0.00 | 0.00 | 0.00 | 0.00 |  |  | range select(-75~100)°C | LOWBATH | [7 empty] | 11 | 2.25 | [7 empty] | MEAN TEMP | 0.09 | 50.09 | 100.12 | 150.11 | 200.04 | 250.23 | 0.00 | 0.00 | [9 empty] | 11 | 2.25
Standard Deviation Mean | 9.2519e-18 | 2.3685e-15 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  |  | range select(100~500)°C | MIDBATH | [7 empty] | 12 | 2.23 | [7 empty] | Standard Deviation Mean | 4.6259e-18 | 2.3685e-15 | 0 | 9.4739e-15 | 0 | 9.4739e-15 | #DIV/0! | #DIV/0! | [9 empty] | 12 | 2.23
[11 empty] | range select(100~1000)°C | HIGHBATH | [7 empty] | 13 | 2.21 | [25 empty] | 13 | 2.21
UUT rdg 1 | 0.1 | 50.1 | 100.1 | [7 empty] | range select(10~70)°C | TEMPCHAMBER | [7 empty] | 14 | 2.2 | [7 empty] | UUT rdg 1 | -0.4 | 49.8 | 99.9 | 149.8 | 199.7 | 249.9 | [11 empty] | 14 | 2.2
UUT rdg 2 | 0.1 | 50.1 | 100.1 | [7 empty] | range select(10~100)%R.H. | HUMIDITYCHAMBER | [7 empty] | 15 | 2.18 | [7 empty] | UUT rdg 2 | -0.4 | 49.8 | 99.9 | 149.8 | 199.7 | 249.9 | [11 empty] | 15 | 2.18
UUT rdg 3 | 0.1 | 50.1 | 100.1 | [7 empty] | range select(0~450)°C | S.PLATE | [7 empty] | 16 | 2.17 | [7 empty] | UUT rdg 3 | -0.4 | 49.8 | 99.9 | 149.8 | 199.7 | 249.9 | [11 empty] | 16 | 2.17
UUT rdg 4 | 0.1 | 50.1 | 100.1 | [16 empty] | 17 | 2.16 | [7 empty] | UUT rdg 4 | -0.4 | 49.8 | 99.9 | 149.8 | 199.7 | 249.9 | [11 empty] | 17 | 2.16
UUT rdg 5 | 0.1 | 50.1 | 100.1 | [7 empty] | DMM-RTD | DMRTD | [7 empty] | 18 | 2.15 | [7 empty] | UUT rdg 5 | -0.4 | 49.8 | 99.9 | 149.8 | 199.7 | 249.9 | [11 empty] | 18 | 2.15
UUT rdg 6 | 0.1 | 50.1 | 100.1 | [7 empty] | DMM 81/2(RTD) | DMRTD8.5 | [7 empty] | 19 | 2.14 | [7 empty] | UUT rdg 6 | -0.4 | 49.8 | 99.9 | 149.8 | 199.7 | 249.9 | [11 empty] | 19 | 2.14
UUT rdg 7 | 0.1 | 50.1 | 100.1 | [7 empty] | DMM-TC | DMTC | [7 empty] | 20 | 2.13 | [7 empty] | UUT rdg 7 | -0.4 | 49.8 | 99.9 | 149.8 | 199.7 | 249.9 | [11 empty] | 20 | 2.13
UUT rdg 8 | 0.1 | 50.1 | 100.1 | [7 empty] | USING GLASS THERMOMETER | GT | [7 empty] | 25 | 2.11 | [7 empty] | UUT rdg 8 | -0.4 | 49.8 | 99.9 | 149.8 | 199.7 | 249.9 | [11 empty] | 25 | 2.11
UUT rdg 9 | 0.1 | 50.1 | 100.1 | [16 empty] | 30 | 2.09 | [7 empty] | UUT rdg 9 | -0.4 | 49.8 | 99.9 | 149.8 | 199.7 | 249.9 | [11 empty] | 30 | 2.09
UUT rdg 10 | 0.1 | 50.1 | 100.1 | [16 empty] | 35 | 2.07 | [7 empty] | UUT rdg 10 | -0.4 | 49.8 | 99.9 | 149.8 | 199.7 | 249.9 | [11 empty] | 35 | 2.07
MEAN TEMP | 0.1 | 50.1 | 100.1 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | [11 empty] | 40 | 2.06 | [7 empty] | MEAN TEMP | -0.4 | 49.8 | 99.9 | 149.8 | 199.7 | 249.9 | 0.0 | 0.0 | [9 empty] | 40 | 2.06
Standard Deviation Mean | 4.6259e-18 | 2.3685e-15 | 4.737e-15 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [11 empty] | 45 | 2.06 | [7 empty] | Standard Deviation Mean | 1.8504e-17 | 2.3685e-15 | 4.737e-15 | 9.4739e-15 | 9.4739e-15 | 9.4739e-15 | #DIV/0! | #DIV/0! | [9 empty] | 45 | 2.06
[20 empty] | 50 | 2.05 | [25 empty] | 50 | 2.05
[20 empty] | 100 | 2.025 | [25 empty] | 100 | 2.025
[20 empty] | 101 | 2 | [25 empty] | 101 | 2
UNCERTAINTY CALCULATION | [28 empty] | UNCERTAINTY CALCULATION
DEGREES OF FREEDOM | Infinity |  | Infinity |  | Infinity |  | Infinity |  | Infinity | Infinity | Infinity | Infinity | Infinity | Infinity | Infinity |  |  | 9 | 9 | Infinity |  | combined Uncertainty | Eff Degrees of freedom | k-factor @ approx 95% CL (from t-distribution table) | expanded uncertainty | round up value ± | REPORT UNCERTAINTY °C |  | DEGREES OF FREEDOM | Infinity |  | Infinity |  | Infinity |  | Infinity |  | Infinity | Infinity | Infinity | Infinity | Infinity |  | Infinity | 9 | 9 | Infinity |  | combined Uncertainty | Eff Degrees of freedom | k-factor @ approx 95% CL (from t-distribution table) | expanded uncertainty | round up value ± | REPORT UNCERTAINTY °C
probability distribution- | Normal |  | Rectangular |  | Normal |  | Rectangular |  | Rectangular | Rectangular | Rectangular | Rectangular | Rectangular | Normal | Rectangular | Rectangular | Rectangular | Normal | Normal | Normal | [8 empty] | probability distribution- | Normal |  | Rectangular |  | Normal |  | Rectangular |  | Rectangular | Rectangular | Rectangular | Rectangular | Rectangular | Normal | Rectangular | Normal | Normal | Rectangular
divisor | 2 |  | 1.7321 |  | 2 |  | 1.7321 |  | 1.7321 | 1.7321 | 1.7321 | 1.7321 | 1.7321 |  |  |  |  | 1 | 1 | 3.4641 |  |  |  |  | °C |  |  |  | divisor | 2 |  | 1.7321 |  | 2 |  | 1.7321 |  | 1.7321 | 1.7321 | 1.7321 | 1.7321 | 1.7321 | 2 | 1.7321 | 1 | 1 | 3.4641 |  |  |  |  | °C
Setpoint | REF SENSOR Uncertainty °C |  | REF SENSOR Accuracy °C |  | REF INDICATOR1 Uncertainty °C |  | REF INDICATOR1 Accuracy °C |  | Chamber Stability °C | Chamber Uniformity °C | Ref Drift(SENSOR) °C | U EFFECT | TC INHOMOGENEITY °C | REF INDR2 Un °C | REF INDR2 Acc °C | Positioning Distance & Focus | Emissivity | Ref Repeatability °C | UUT Repeatability °C | Resolution °C | [8 empty] | Setpoint | REF sensor Uncertainty °C |  | REF Accuracy °C |  | Ref Indicator 1 Uncertainty °C |  | Ref Indicator 1 Accuracy °C |  | Chamber Stability °C | Chamber Uniformity °C | Ref Drift °C | U EFFECT | tc inhomogeneity °C | REF INDR2 Un °C | REF INDR2 Acc °C | Ref Repeatability °C | UUT Repeatability °C | Resolution °C
°C | [28 empty] | °C
0.14 | 0.018 | 0.009 | 0.2 | 0.11547 | 0.05 | 0.025 | 0.1 | 0.057735 | 0.057735 | 0.046188 | 0.10392 | 0.057735 | 0 | 0 | 0 | 0 | 0 | 9.2519e-18 | 4.6259e-18 | 0.1 | 0.028868 | 0.19444 | 1.6524e+66 | 2 | 0.38888 | 0.38888 | 0.38888 |  | 0.09 | 0.018 | 0.009 | 0.2 | 0.11547 | 0.05 | 0.025 | 0.1 | 0.057735 | 0.057735 | 0.046188 | 0.10392 | 0.057735 | 0 | 0 | 0 | 4.6259e-18 | 1.8504e-17 | 0.1 | 0.028868 | 0.19444 | 1.093e+65 | 2 | 0.38888 | 0.38888 | 0.38888
50.35 | 0.018 | 0.009 | 0.2 | 0.11547 | 0.05 | 0.025 | 0.1 | 0.057735 | 0.057735 | 0.046188 | 0.10392 | 0.057735 | 0 | 0 | 0 | 0 | 0 | 2.3685e-15 | 2.3685e-15 | 0.1 | 0.028868 | 0.19444 | 2.0439e+56 | 2 | 0.38888 | 0.38888 | 0.38888 |  | 50.09 | 0.018 | 0.009 | 0.2 | 0.11547 | 0.05 | 0.025 | 0.1 | 0.057735 | 0.057735 | 0.046188 | 0.10392 | 0.057735 | 0 | 0 | 0 | 2.3685e-15 | 2.3685e-15 | 0.1 | 0.028868 | 0.19444 | 2.0439e+56 | 2 | 0.38888 | 0.38888 | 0.38888
100.48 | 0.018 | 0.009 | 0.2 | 0.11547 | 0.05 | 0.025 | 0.1 | 0.057735 | 0.057735 | 0.086603 | 0.10392 | 0.057735 | 0 | 0 | 0 | 0 | 0 | 0 | 4.737e-15 | 0.1 | 0.028868 | 0.20778 | 3.3317e+55 | 2 | 0.41556 | 0.41556 | 0.41556 |  | 100.12 | 0.018 | 0.009 | 0.2 | 0.11547 | 0.05 | 0.025 | 0.1 | 0.057735 | 0.057735 | 0.086603 | 0.10392 | 0.057735 | 0 | 0 | 0 | 0 | 4.737e-15 | 0.1 | 0.028868 | 0.20778 | 3.3317e+55 | 2 | 0.41556 | 0.41556 | 0.41556
0.00 | FALSE | 0 | FALSE | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.057735 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0.1 | 0.028868 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  | 150.11 | FALSE | 0 | FALSE | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.057735 | 0 | 0 | 0 | 9.4739e-15 | 9.4739e-15 | 0.1 | 0.028868 | 0.06455 | 9.6978e+51 | 2 | 0.1291 | 0.1291 | 0.1291
0.00 | FALSE | 0 | FALSE | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.057735 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0.1 | 0.028868 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  | 200.04 | FALSE | 0 | FALSE | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.057735 | 0 | 0 | 0 | 0 | 9.4739e-15 | 0.1 | 0.028868 | 0.06455 | 1.9396e+52 | 2 | 0.1291 | 0.1291 | 0.1291
0.00 | FALSE | 0 | FALSE | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.057735 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0.1 | 0.028868 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  | 250.23 | FALSE | 0 | FALSE | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.057735 | 0 | 0 | 0 | 9.4739e-15 | 9.4739e-15 | 0.1 | 0.028868 | 0.06455 | 9.6978e+51 | 2 | 0.1291 | 0.1291 | 0.1291
0.00 | FALSE | 0 | FALSE | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.057735 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0.1 | 0.028868 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  | 0.00 | FALSE | 0 | FALSE | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.057735 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0.1 | 0.028868 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
0.00 | FALSE | 0 | FALSE | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.057735 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0.1 | 0.028868 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  | 0.00 | FALSE | 0 | FALSE | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.057735 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0.1 | 0.028868 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
HUMIDITY | THMHUM
UNIT % r.h.
SET POINT % r.h. | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
Ref.inst.rdg 1 | 35.2 | 55.8 | 86.4
Ref.inst.rdg 2 | 35.2 | 55.8 | 86.4
Ref.inst.rdg 3 | 35.2 | 55.8 | 86.4
Ref.inst.rdg 4 | 35.2 | 55.8 | 86.4
Ref.inst.rdg 5 | 35.2 | 55.8 | 86.4
Ref.inst.rdg 6 | 35.2 | 55.8 | 86.4
Ref.inst.rdg 7 | 35.2 | 55.8 | 86.4
Ref.inst.rdg 8 | 35.2 | 55.8 | 86.4
Ref.inst.rdg 9 | 35.2 | 55.8 | 86.4
Ref.inst.rdg 10 | 35.2 | 55.8 | 86.4
MEAN TEMP | 35.2 | 55.8 | 86.4 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0
Standard Deviation Mean | 2.3685e-15 | 0 | 4.737e-15 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
UUT rdg 1 | 35 | 55 | 85
UUT rdg 2 | 35 | 55 | 85
UUT rdg 3 | 35 | 55 | 85
UUT rdg 4 | 35 | 55 | 85
UUT rdg 5 | 35 | 55 | 85
UUT rdg 6 | 35 | 55 | 85
UUT rdg 7 | 35 | 55 | 85
UUT rdg 8 | 35 | 55 | 85
UUT rdg 9 | 35 | 55 | 85
UUT rdg 10 | 35 | 55 | 85
MEAN TEMP | 35 | 55 | 85 | 0 | 0 | 0 | 0 | 0
Standard Deviation Mean | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
UNCERTAINTY CALCULATION
DEGREES OF FREEDOM | Infinity |  | Infinity |  | Infinity |  | Infinity |  | Infinity | Infinity | Infinity | Infinity | Infinity |  |  |  |  | 9 | 9 | Infinity |  | combined Uncertainty | Eff Degrees of freedom | k-factor @ approx 95% CL (from t-distribution table) | expanded uncertainty | round up value ± | REPORT UNCERTAINTY %r.h.
probability distribution- | Normal |  | Rectangular |  | Normal |  | Rectangular |  | Rectangular | Rectangular | Rectangular | Rectangular | Rectangular |  |  |  |  | Normal | Normal | Rectangular
divisor | 2 |  | 1.7321 |  | 2 |  | 1.7321 |  | 1.7321 | 1.7321 | 1.7321 | 1.7321 | 1.7321 |  |  |  |  | 1 | 1 | 3.4641 |  |  |  |  | %r.h.
Setpoint | REF Uncertainty %r.h. |  | REF Accuracy %r.h. |  | Indicator Uncertainty %r.h. |  | Indicator Accuracy %r.h. |  | Chamber Stability %r.h. | Chamber Uniformity %r.h. | Ref Drift %r.h. | Temperature effect | hysterisesh %r.h. |  |  |  |  | Ref Repeatability %r.h. | UUT Repeatability %r.h. | Resolution %r.h.
%r.h.
35.2 | 1.5 | 0.75 | 1.325 | 0.76499 | 0 | 0 | 0 | 0 | 0.57735 | 0.86603 | 0.75056 | 0.63509 | 0 |  |  |  |  | 2.3685e-15 | 0 | 0 | 0 | 1.7882 | 2.9245e+60 | 2 | 3.5764 | 3.5764 | 3.5764
55.8 | 1.5 | 0.75 | 1.775 | 1.0248 | 0 | 0 | 0 | 0 | 0.57735 | 0.86603 | 0.75056 | 0.63509 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 1.9138 | #DIV/0! | 2 | 3.8276 | 3.8276 | 3.8276
86.4 | 1.5 | 0.75 | 2 | 1.1547 | 0 | 0 | 0 | 0 | 0.57735 | 0.86603 | 0.75056 | 0.63509 | 0 |  |  |  |  | 4.737e-15 | 0 | 0 | 0 | 1.9864 | 2.7831e+59 | 2 | 3.9728 | 3.9728 | 3.9728
0.0 | 1.5 | 0.75 | 0.5 | 0.28868 | 0 | 0 | 0 | 0 | 0.57735 | 0.86603 | 0.75056 | 0.63509 | 0 |  |  |  |  | #DIV/0! | #DIV/0! | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
0.0 | 1.5 | 0.75 | 0.5 | 0.28868 | 0 | 0 | 0 | 0 | 0.57735 | 0.86603 | 0.75056 | 0.63509 | 0 |  |  |  |  | #DIV/0! | #DIV/0! | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
0.0 | 1.5 | 0.75 | 0.5 | 0.28868 | 0 | 0 | 0 | 0 | 0.57735 | 0.86603 | 0.75056 | 0.63509 | 0 |  |  |  |  | #DIV/0! | #DIV/0! | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
0.0 | 1.5 | 0.75 | 0.5 | 0.28868 | 0 | 0 | 0 | 0 | 0.57735 | 0.86603 | 0.75056 | 0.63509 | 0 |  |  |  |  | #DIV/0! | #DIV/0! | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
0.0 | 1.5 | 0.75 | 0.5 | 0.28868 | 0 | 0 | 0 | 0 | 0.57735 | 0.86603 | 0.75056 | 0.63509 | 0 |  |  |  |  | #DIV/0! | #DIV/0! | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Note on G1: DO N'T TYPE THIS ROW
Note on C8: User: select reference sensor like rtd ,thermocouple
Note on D8: User: select reference sensor like rtd ,thermocouple
Note on E8: User: select reference sensor like rtd ,thermocouple
Note on F8: User: select reference sensor like rtd ,thermocouple
Note on G8: User: select reference sensor like rtd ,thermocouple
Note on H8: User: select reference sensor like rtd ,thermocouple
Note on I8: User: select reference sensor like rtd ,thermocouple
Note on J8: User: select reference sensor like rtd ,thermocouple
Note on C9: User: select indicator like digital multimeter ohms mode(DMRTD) or millivolt mode(DMTC)
Note on D9: User: select indicator like digital multimeter ohms mode(DMRTD) or millivolt mode(DMTC)
Note on E9: User: select indicator like digital multimeter ohms mode(DMRTD) or millivolt mode(DMTC)
Note on F9: User: select indicator like digital multimeter ohms mode(DMRTD) or millivolt mode(DMTC)
Note on G9: User: select indicator like digital multimeter ohms mode(DMRTD) or millivolt mode(DMTC)
Note on H9: User: select indicator like digital multimeter ohms mode(DMRTD) or millivolt mode(DMTC)
Note on I9: User: select indicator like digital multimeter ohms mode(DMRTD) or millivolt mode(DMTC)
Note on J9: User: select indicator like digital multimeter ohms mode(DMRTD) or millivolt mode(DMTC)
Note on C10: User: select heat source using temperature bath or Infrared
Note on D10: User: select heat source using temperature bath or Infrared
Note on E10: User: select heat source using temperature bath or Infrared
Note on F10: User: select heat source using temperature bath or Infrared
Note on G10: User: select heat source using temperature bath or Infrared
Note on H10: User: select heat source using temperature bath or Infrared
Note on I10: User: select heat source using temperature bath or Infrared
Note on J10: User: select heat source using temperature bath or Infrared
Note on B11: User: USING MASTER OF THERMOCOUPLE AND GLASS THERMOMETER
Note on C11: User: select indicator like digital multimeter ohms mode(DMRTD) or millivolt mode(DMTC),Glass ther
Note on D11: User: select indicator like digital multimeter ohms mode(DMRTD) or millivolt mode(DMTC),Glass ther
Note on E11: User: select indicator like digital multimeter ohms mode(DMRTD) or millivolt mode(DMTC),Glass ther
Note on F11: User: select indicator like digital multimeter ohms mode(DMRTD) or millivolt mode(DMTC),Glass ther
Note on G11: User: select indicator like digital multimeter ohms mode(DMRTD) or millivolt mode(DMTC),Glass ther
Note on H11: User: select indicator like digital multimeter ohms mode(DMRTD) or millivolt mode(DMTC),Glass ther
Note on I11: User: select indicator like digital multimeter ohms mode(DMRTD) or millivolt mode(DMTC),Glass ther
Note on J11: User: select indicator like digital multimeter ohms mode(DMRTD) or millivolt mode(DMTC),Glass ther
Note on AI50: bs-electrical: /1 IS SENSITIVITY CO EFFICIENT
Note on F100: bs-electrical: /1 IS SENSITIVITY CO EFFICIENT

## Sheet1 (hidden)
|  |  | SPRT | (SPRT) | [6 empty] | PRT | (PRT)
LOGIC | S.NO | TEMPERATURE | UNCERTAINTY | ACCURACY | DRIFT | OTHER |  | LOGIC | S.NO | TEST POINT | UNCERTAINTY | ACCURACY | DRIFT | OTHER
3 | 1 | 500 | 0.025 | 0.01 | 0.027 |  |  | 3 | 1 | 500 | 0.018 | 0.2 | 0.18
3 | 2 | 250 | 0.025 | 0.01 | 0.027 |  |  | 3 | 2 | 250 | 0.018 | 0.2 | 0.18
2 | 3 | 100 | 0.025 | 0.01 | 0.027 |  |  | 2 | 3 | 100 | 0.018 | 0.2 | 0.18
4 | 4 | 0 | 0.025 | 0.01 | 0.027 |  |  | 4 | 4 | 0 | 0.018 | 0.2 | 0.18
4 | 5 | -30 | 0.025 | 0.01 | 0.027 |  |  | 4
4 | 6 | -75 | 0.025 | 0.01 | 0.027 |  |  | 4
4 | [7 empty] | 4
4 | [7 empty] | 4
4 | [7 empty] | 4
4 | [7 empty] | 4
4 | [7 empty] | 4
4 | [7 empty] | 4
4 | [7 empty] | 4
4 | [7 empty] | 4
|  |  | 8.5 DMM FOR OHMS | (DMRTD8.5) | [6 empty] | 6.5 DMM FOR OHMS |  | (DMRTD)
LOGIC | S.NO | TEST POINT | UNCERTAINTY | ACCURACY | DRIFT | OTHER |  | LOGIC | S.NO | TEST POINT | UNCERTAINTY | ACCURACY | DRIFT | OTHER
3 | 1 | 500 | 0.008 | 0.017 |  |  |  | 3 | 1 | 500 | 0.05 | 0.1
3 | 2 | 250 | 0.008 | 0.017 |  |  |  | 3 | 2 | 250 | 0.05 | 0.1
2 | 3 | 100 | 0.004 | 0.009 |  |  |  | 2 | 3 | 100 | 0.05 | 0.1
4 | 4 | 0 | 0.004 | 0.009 |  |  |  | 4 | 4 | 0 | 0.05 | 0.1
4 | 5 | -30 | 0.004 | 0.009 |  |  |  | 4 | 5 | -30 | 0.05 | 0.1
4 | 6 | -75 | 0.004 | 0.009 |  |  |  | 4 | 6 | -75 | 0.05 | 0.1
4 | [7 empty] | 4
4 | [7 empty] | 4
4 | [7 empty] | 4
4 | [7 empty] | 4
4 | [7 empty] | 4
4 | [7 empty] | 4
4 | [7 empty] | 4
4 | [7 empty] | 4
|  |  | TC(S-TYPE) | (STC) | [5 empty] | DMMTC | (DMTC)
LOGIC | S.NO | TEST POINT | UNCERTAINTY | ACCURACY | DRIFT | INHOMOGENEITY |  | LOGIC | S.NO | TEST POINT | UNCERTAINTY | ACCURACY | OTHER | OTHER
5 | 1 | 1000 | 0.6 | 0.5 | 0.5 | 0.2 |  | 4 | 1 | 1000 | 0.3 | 0.4
4 | 2 | 750 | 0.6 | 0.5 | 0.1 | 0.2 |  | 4 | 2 | 800 | 0.3 | 0.4
4 | 3 | 500 | 0.6 | 0.5 | 0.1 | 0.1 |  | 4 | 3 | 500 | 0.3 | 0.4
6 | 4 | 250 | 0.4 | 0.5 | 0.1 | 0.1 |  | 5 | 4 | 300 | 0.3 | 0.4
6 | 5 | 50 | 0.4 | 0.5 | 0.1 | 0.1 |  | 5 | 5 | 0 | 0.3 | 0.4
6 | 6 | 0 | 0.4 | 0.5 | 0.1 | 0.1 |  | 5 | 6 | -75 | 0.3 | 0.4
6 | [7 empty] | 5
6 | [7 empty] | 5
6 | [7 empty] | 5
6 | [7 empty] | 5
6 | [7 empty] | 5
6 | [7 empty] | 5
6 | [7 empty] | 5
6 | [7 empty] | 5
|  |  | THM | (THMTEMP) | [6 empty] | THM | (THMHUM) |  |  | TEMP&HUMITY
BOTH LOGIC
LOGIC | S.NO | TEST POINT | UNCERTAINTY | ACCURACY | DRIFT | OTHER |  | LOGIC | S.NO | TEST POINT | UNCERTAINTY | ACCURACY | DRIFT | OTHER | LOGIC
4 | 1 | 80 | 0.2 | 0.2 | 0.2 |  |  | 4 | 1 | 100 | 1.5 | 2 | 1.3 |  | 3
1 | 2 | 50 | 0.2 | 0.2 | 0.2 |  |  | 3 | 2 | 85 | 1.5 | 1.775 | 1.3 |  | 2
#N/A | 3 | 20 | 0.2 | 0.2 | 0.2 |  |  | #N/A | 3 | 55 | 1.5 | 1.325 | 1.3 |  | 1
5 | 4 | 10 | 0.2 | 0.2 | 0.2 |  |  | 5 | 4 | 35 | 1.5 | 1.025 | 1.3 |  | 5
5 | 5 | 0 | 0.2 | 0.2 | 0.2 |  |  | 5 | 5 | 0 | 1.5 | 0.5 | 1.3 |  | 5
5 | [7 empty] | 5 | [6 empty] | 5
5 | [7 empty] | 5 | [6 empty] | 5
5 | [7 empty] | 5 | [6 empty] | 5
5 | [7 empty] | 5 | [6 empty] | 5
5 | [7 empty] | 5 | [6 empty] | 5
5 | [7 empty] | 5 | [6 empty] | 5
5 | [7 empty] | 5 | [6 empty] | 5
5 | [7 empty] | 5 | [6 empty] | 5
5 | [7 empty] | 5 | [6 empty] | 5
|  |  | IR | (IR) | [7 empty] | GT | (GT)
LOGIC | S.NO | TEST POINT | Positioning Distance & Focus | Emissivity | DRIFT | STABILITY | UNIFORMITY |  | LOGIC | S.NO | TEST POINT | ELC ERROR | RTD UNCETAINTY
5 | 1 | 1000 | 0.5 | 6 | 0.1 | 4 | 4.5 |  | 3 | 1 | 500 | 0.22 | 0.08
5 | 2 | 800 | 0.5 | 4.8 | 0.1 | 4 | 4.5 |  | 3 | 2 | 300 | 0.28 | 0.08
5 | 3 | 600 | 0.5 | 3.6 | 0.1 | 4 | 4.5 |  | 2 | 3 | 100 | 0.28 | 0.08
6 | 4 | 300 | 0.5 | 2.7 | 0.1 | 4 | 4.5 |  | 4 | 4 | 0 | 0.32 | 0.08
6 | 5 | 200 | 0.5 | 1.3 | 0.1 | 4 | 4.5 |  | 4 | 5 | -75 | 0.32 | 0.08
6 | 6 | 0 | 0.5 | 1.3 | 0.1 | 4 | 4.5 |  | 4
6 | [8 empty] | 4
6 | [8 empty] | 4
6 | [8 empty] | 4
6 | [8 empty] | 4
6 | [8 empty] | 4
6 | [8 empty] | 4
6 | [8 empty] | 4
6 | [8 empty] | 4

## Sheet2 (hidden)
HEAT SOURCE
|  | LOW TEMPERATURE BATH |  |  | (LOWBATH) |  |  |  | MID RANGE BATH |  | (MIDBATH)
LOGIC | S.NO | TEST POINT | STABILITY | UNIFORMITY | OTHER | OTHER | LOGIC | S.NO | TEST POINT | STABILITY | UNIFORMITY | OTHER | OTHER
1 | 1 | 100 | 0.1 | 0.08 |  |  | 3 | 1 | 500 | 0.1 | 0.15
1 | 2 | 0 | 0.08 | 0.08 |  |  | 3 | 2 | 300 | 0.1 | 0.15
#N/A | 3 | -30 | 0.12 | 0.1 |  |  | 2 | 3 | 100 | 0.1 | 0.15
2 | 4 | -75 | 0.12 | 0.1 |  |  | 4 | 4 | 0 | 0.1 | 0.15
2 | 5 | [5 empty] | 4 | 5
2 | 6 | [5 empty] | 4 | 6
2 | 7 | [5 empty] | 4 | 7
2 | [6 empty] | 4
2 | [6 empty] | 4
2 | [6 empty] | 4
2 | [6 empty] | 4
2 | [6 empty] | 4
2 | [6 empty] | 4
2 | [6 empty] | 4
100
|  | FURNACE HIGH TEMPERATURE BATH |  |  |  | (HIGHBATH) | 4 |  |  | SURFACE PLATE |  | (S.PLATE)
LOGIC | S.NO | TEST POINT | STABILITY | UNIFORMITY | OTHER | OTHER | LOGIC | S.NO | TEST POINT | STABILITY | UNIFORMITY | DRIFT | OTHER
5 | 1 | 1000 | 0.5 | 0.7 |  |  | 5 | 1 | 450 | 3.1 | 3.5
5 | 2 | 800 | 0.5 | 0.7 |  |  | 4 | 2 | 400 | 3.1 | 3.5
4 | 3 | 500 | 0.5 | 0.7 |  |  | 3 | 3 | 250 | 1.6 | 2
6 | 4 | 300 | 0.5 | 0.7 |  |  | 6 | 4 | 100 | 1.6 | 2
6 | 5 | 100 | 0.5 | 0.7 |  |  | 6 | 5 | 50 | 1.6 | 2
6 | 6 | 0 | 0.5 | 0.7 |  |  | 6 | 6 | 0 | 1.6 | 2
6 | 7 | [5 empty] | 6 | 7
6 | [6 empty] | 6
6 | [6 empty] | 6
6 | [6 empty] | 6
6 | [6 empty] | 6
6 | [6 empty] | 6
6 | [6 empty] | 6
6 | [6 empty] | 6
|  | HUMIDITY CHAMBER(TEMP) |  |  | (TEMPCHAMBER) |  |  |  | HUMIDITY CHAMBER(HUM) |  |  | (HUMIDITYCHAMBER)
LOGIC | S.NO | TEST POINT | STABILITY | UNIFORMITY | OTHER | OTHER | LOGIC | S.NO | TEST POINT | STABILITY | UNIFORMITY | OTHER | OTHER
4 | 1 | 70 | 0.15 | 0.2 |  |  | 5 | 1 | 100 | 1 | 1.5
1 | 2 | 40 | 0.15 | 0.2 |  |  | 2 | 2 | 80 | 1 | 1.5
#N/A | 3 | 20 | 0.15 | 0.2 |  |  | #N/A | 3 | 50 | 1 | 1.5
4 | 4 | 10 | 0.15 | 0.2 |  |  | 5 | 4 | 30 | 1 | 1.5
4 | 5 | [5 empty] | 5 | 5 | 10 | 1 | 1.5
4 | 6 | [5 empty] | 5 | 6
4 | 7 | [5 empty] | 5 | 7
4 | [6 empty] | 5
4 | [6 empty] | 5
4 | [6 empty] | 5
4 | [6 empty] | 5
4 | [6 empty] | 5
4 | [6 empty] | 5
4 | [6 empty] | 5

## Thermometer Raw Data
|  | RAW DATA | [14 empty] | Ambient Temperature |  |  |  |  | : | (21 ± 2) °C
Certificate no. | TS 130353
RAW DATA | [23 empty] | RAW DATA
TEMPERATURE | [12 empty] | J TYPE |  | degree | Coverage factor | [7 empty] | TEMPERATURE | [14 empty] | degree | Coverage factor
UNIT °C | [12 empty] | K TYPE |  | of freedom | k factor | [7 empty] | UNIT °C | [14 empty] | of freedom | k factor
MEASURE MODE | K TYPE | [11 empty] | T TYPE
SOURCE MODE | [12 empty] | R TYPE
SET POINT °C | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |  |  |  |  | S TYPE | [10 empty] | SET POINT °C | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
Ref.inst.rdg 1 | -100 | 0 | 500 | 1000 | 1200 | [7 empty] | B TYPE |  | 1 | 13.97 | [7 empty] | Ref.inst.rdg 1 | -75.03 | 150.32 | 850.68 | [11 empty] | 1 | 13.97
Ref.inst.rdg 2 | -100 | 0 | 500 | 1000 | 1200 | [7 empty] | N TYPE |  | 2 | 4.53 | [7 empty] | Ref.inst.rdg 2 | -75.03 | 150.32 | 850.68 | [11 empty] | 2 | 4.53
Ref.inst.rdg 3 | -100 | 0 | 500 | 1000 | 1200 | [7 empty] | E TYPE |  | 3 | 3.31 | [7 empty] | Ref.inst.rdg 3 | -75.03 | 150.32 | 850.68 | [11 empty] | 3 | 3.31
Ref.inst.rdg 4 | -100 | 0 | 500 | 1000 | 1200 | [7 empty] | PRT(PT 100-385) |  | 4 | 2.87 | [7 empty] | Ref.inst.rdg 4 | -75.03 | 150.32 | 850.68 | [11 empty] | 4 | 2.87
Ref.inst.rdg 5 | -100 | 0 | 500 | 1000 | 1200 | [9 empty] | 5 | 2.65 | [7 empty] | Ref.inst.rdg 5 | -75.03 | 150.32 | 850.68 | [11 empty] | 5 | 2.65
Ref.inst.rdg 6 | -100 | 0 | 500 | 1000 | 1200 | [9 empty] | 6 | 2.52 | [7 empty] | Ref.inst.rdg 6 | -75.03 | 150.32 | 850.68 | [11 empty] | 6 | 2.52
Ref.inst.rdg 7 | -100 | 0 | 500 | 1000 | 1200 | [9 empty] | 7 | 2.43 | [7 empty] | Ref.inst.rdg 7 | -75.03 | 150.32 | 850.68 | [11 empty] | 7 | 2.43
Ref.inst.rdg 8 | -100 | 0 | 500 | 1000 | 1200 | [9 empty] | 8 | 2.37 | [7 empty] | Ref.inst.rdg 8 | -75.03 | 150.32 | 850.68 | [11 empty] | 8 | 2.37
Ref.inst.rdg 9 | -100 | 0 | 500 | 1000 | 1200 | [9 empty] | 9 | 2.32 | [7 empty] | Ref.inst.rdg 9 | -75.03 | 150.32 | 850.68 | [11 empty] | 9 | 2.32
Ref.inst.rdg 10 | -100 | 0 | 500 | 1000 | 1200 | [9 empty] | 10 | 2.28 | [7 empty] | Ref.inst.rdg 10 | -75.03 | 150.32 | 850.68 | [11 empty] | 10 | 2.28
MEAN TEMP | -100.00 | 0.00 | 500.00 | 1,000.00 | 1,200.00 | 0.00 | 0.00 | 0.00 | [6 empty] | 11 | 2.25 | [7 empty] | MEAN TEMP | -75.03 | 150.32 | 850.68 | 0.00 | 0.00 | 0.00 | 0.00 | 0.00 | [6 empty] | 11 | 2.25
Standard Deviation Mean | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | [6 empty] | 12 | 2.23 | [7 empty] | Standard Deviation Mean | 4.737e-15 | 9.4739e-15 | 3.7896e-14 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [6 empty] | 12 | 2.23
[15 empty] | 13 | 2.21 | [22 empty] | 13 | 2.21
UUT rdg 1 | -100.2 | -0.2 | 499.8 | 999.9 | 1200 | [9 empty] | 14 | 2.2 | [7 empty] | UUT rdg 1 | -74.8 | 150.1 | 848.3 | [11 empty] | 14 | 2.2
UUT rdg 2 | -100.2 | -0.2 | 499.8 | 999.9 | 1200 | [9 empty] | 15 | 2.18 | [7 empty] | UUT rdg 2 | -74.8 | 150.1 | 848.3 | [11 empty] | 15 | 2.18
UUT rdg 3 | -100.2 | -0.2 | 499.8 | 999.9 | 1200 | [9 empty] | 16 | 2.17 | [7 empty] | UUT rdg 3 | -74.8 | 150.1 | 848.3 | [11 empty] | 16 | 2.17
UUT rdg 4 | -100.2 | -0.2 | 499.8 | 999.9 | 1200 | [9 empty] | 17 | 2.16 | [7 empty] | UUT rdg 4 | -74.8 | 150.1 | 848.3 | [11 empty] | 17 | 2.16
UUT rdg 5 | -100.2 | -0.2 | 499.8 | 999.9 | 1200 | [9 empty] | 18 | 2.15 | [7 empty] | UUT rdg 5 | -74.8 | 150.1 | 848.3 | [11 empty] | 18 | 2.15
UUT rdg 6 | -100.2 | -0.2 | 499.8 | 999.9 | 1200 | [9 empty] | 19 | 2.14 | [7 empty] | UUT rdg 6 | -74.8 | 150.1 | 848.3 | [11 empty] | 19 | 2.14
UUT rdg 7 | -100.2 | -0.2 | 499.8 | 999.9 | 1200 | [9 empty] | 20 | 2.13 | [7 empty] | UUT rdg 7 | -74.8 | 150.1 | 848.3 | [11 empty] | 20 | 2.13
UUT rdg 8 | -100.2 | -0.2 | 499.8 | 999.9 | 1200 | [9 empty] | 25 | 2.11 | [7 empty] | UUT rdg 8 | -74.8 | 150.1 | 848.3 | [11 empty] | 25 | 2.11
UUT rdg 9 | -100.2 | -0.2 | 499.8 | 999.9 | 1200 | [9 empty] | 30 | 2.09 | [7 empty] | UUT rdg 9 | -74.8 | 150.1 | 848.3 | [11 empty] | 30 | 2.09
UUT rdg 10 | -100.2 | -0.2 | 499.8 | 999.9 | 1200 | [9 empty] | 35 | 2.07 | [7 empty] | UUT rdg 10 | -74.8 | 150.1 | 848.3 | [11 empty] | 35 | 2.07
MEAN TEMP | -100.2 | -0.2 | 499.8 | 999.9 | 1,200.0 | 0.0 | 0.0 | 0.0 | [6 empty] | 40 | 2.06 | [7 empty] | MEAN TEMP | -74.8 | 150.1 | 848.3 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | [6 empty] | 40 | 2.06
Standard Deviation Mean | 4.737e-15 | 9.2519e-18 | 1.8948e-14 | 3.7896e-14 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | [6 empty] | 45 | 2.06 |  | 2 | [5 empty] | Standard Deviation Mean | 4.737e-15 | 9.4739e-15 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [6 empty] | 45 | 2.06
[15 empty] | 50 | 2.05 | [22 empty] | 50 | 2.05
[15 empty] | 100 | 2.025 | [22 empty] | 100 | 2.025
[15 empty] | 101 | 2 | [22 empty] | 101 | 2
UNCERTAINTY CALCULATION | [23 empty] | UNCERTAINTY CALCULATION
DEGREES OF FREEDOM | Infinity |  | Infinity |  | Infinity |  | Infinity |  | Infinity | Infinity | Infinity | Infinity | 9 | 9 | Infinity |  | combined Uncertainty | Eff Degrees of freedom | k-factor @ approx 95% CL (from t-distribution table) | expanded uncertainty | round up value ± | REPORT UNCERTAINTY °C |  | DEGREES OF FREEDOM | Infinity |  | Infinity |  | Infinity |  | Infinity |  | Infinity | Infinity |  | Infinity | 9 | 9 | Infinity |  | combined Uncertainty | Eff Degrees of freedom | k-factor @ approx 95% CL (from t-distribution table) | expanded uncertainty | round up value ± | REPORT UNCERTAINTY °C
probability distribution- | Normal |  | Rectangular |  | Normal |  | Rectangular |  | normal | Rectangular | normal | Rectangular | Normal | Normal | Normal | [8 empty] | probability distribution- | Normal |  | Rectangular |  | Normal |  | Rectangular |  | Rectangular | Rectangular | Normal | Rectangular | Normal | Normal | Rectangular
divisor | 2 |  | 1.7321 |  | 2 |  | 1.7321 |  | 2 | 1.7321 | 2 | 1.7321 | 1 | 1 | 3.4641 |  |  |  |  | °C |  |  |  | divisor | 2 |  | 1.7321 |  | 2 |  | 1.7321 |  | 1.7321 | 1.7321 | 2 | 1.7321 | 1 | 1 | 3.4641 |  |  |  |  | °C
Setpoint | REF Uncertainty °C |  | REF Accuracy °C |  | REF T/C Uncertainty °C |  | REF T/C Accuracy °C |  | REF UC(SOR) °C | REF accu(SOR) °C | REF t/c UC(SOR) °C | REF t/c accu(SOR) °C | Ref Repeatability °C | UUT Repeatability °C | Resolution °C | [8 empty] | Setpoint | REF Uncertainty °C |  | REF Accuracy °C |  | T/C Uncertainty °C |  | T/C Accuracy °C |  | REF(SOU)UC °C | REF(SOUac °C | REF t/c UC(SOR) °C | REF t/c accu(SOR) °C | Ref Repeatability °C | UUT Repeatability °C | Resolution °C
°C | [23 empty] | °C
-100.00 | 0.05 | 0.025 | 0.16 | 0.092376 | 0.4 | 0.2 | 0.08 | 0.046188 | 0 | 0 | 0 | 0 | 0 | 4.737e-15 | 0.1 | 0.028868 | 0.22831 | 4.8567e+55 | 2 | 0.45662 | 0.45662 | 0.45662 |  | -75.03 | 0.05 | 0.025 | 0.16 | 0.092376 | 0.4 | 0.2 | 0.08 | 0.046188 | 0 | 0 | 0 | 0 | 4.737e-15 | 4.737e-15 | 0.1 | 0.028868 | 0.22831 | 2.4283e+55 | 2 | 0.45662 | 0.45662 | 0.45662
0.00 | 0.05 | 0.025 | 0.16 | 0.092376 | 0.4 | 0.2 | 0.08 | 0.046188 | 0 | 0 | 0 | 0 | 0 | 9.2519e-18 | 0.1 | 0.028868 | 0.22831 | 3.3375e+66 | 2 | 0.45662 | 0.45662 | 0.45662 |  | 150.32 | 0.05 | 0.025 | 0.16 | 0.092376 | 0.4 | 0.2 | 0.08 | 0.046188 | 0 | 0 | 0 | 0 | 9.4739e-15 | 9.4739e-15 | 0.1 | 0.028868 | 0.22831 | 1.5177e+54 | 2 | 0.45662 | 0.45662 | 0.45662
500.00 | 0.05 | 0.025 | 0.26 | 0.15011 | 0.4 | 0.2 | 0.08 | 0.046188 | 0 | 0 | 0 | 0 | 0 | 1.8948e-14 | 0.1 | 0.028868 | 0.25715 | 3.0531e+53 | 2 | 0.5143 | 0.5143 | 0.5143 |  | 850.68 | 0.05 | 0.025 | 0.26 | 0.15011 | 0.4 | 0.2 | 0.08 | 0.046188 | 0 | 0 | 0 | 0 | 3.7896e-14 | 0 | 0.1 | 0.028868 | 0.25715 | 1.9082e+52 | 2 | 0.5143 | 0.5143 | 0.5143
1,000.00 | 0.05 | 0.025 | 0.4 | 0.23094 | 0.4 | 0.2 | 0.08 | 0.046188 | 0 | 0 | 0 | 0 | 0 | 3.7896e-14 | 0.1 | 0.028868 | 0.31133 | 4.0998e+52 | 2 | 0.62266 | 0.62266 | 0.62266 |  | 0.00 | 0.05 | 0.025 | 0.4 | 0.23094 | 0.4 | 0.2 | 0.08 | 0.046188 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0.1 | 0.028868 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1,200.00 | 0.05 | 0.025 | 0.4 | 0.23094 | 0.4 | 0.2 | 0.08 | 0.046188 | 0 | 0 | 0 | 0 | 0 | 0 | 0.1 | 0.028868 | 0.31133 | #DIV/0! | 2 | 0.62266 | 0.62266 | 0.62266 |  | 0.00 | 0.05 | 0.025 | 0.4 | 0.23094 | 0.4 | 0.2 | 0.08 | 0.046188 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0.1 | 0.028868 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
0.00 | 0.05 | 0.025 | 0.16 | 0.092376 | 0.4 | 0.2 | 0.08 | 0.046188 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0.1 | 0.028868 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  | 0.00 | 0.05 | 0.025 | 0.16 | 0.092376 | 0.4 | 0.2 | 0.08 | 0.046188 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0.1 | 0.028868 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
0.00 | 0.05 | 0.025 | 0.16 | 0.092376 | 0.4 | 0.2 | 0.08 | 0.046188 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0.1 | 0.028868 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  | 0.00 | 0.05 | 0.025 | 0.16 | 0.092376 | 0.4 | 0.2 | 0.08 | 0.046188 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0.1 | 0.028868 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
0.00 | 0.05 | 0.025 | 0.16 | 0.092376 | 0.4 | 0.2 | 0.08 | 0.046188 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0.1 | 0.028868 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  | 0.00 | 0.05 | 0.025 | 0.16 | 0.092376 | 0.4 | 0.2 | 0.08 | 0.046188 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0.1 | 0.028868 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
HUMIDITY | THMHUM
UNIT % r.h.
SET POINT % r.h. | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
Ref.inst.rdg 1 | 35.2 | 55.8 | 86.4
Ref.inst.rdg 2 | 35.2 | 55.8 | 86.4
Ref.inst.rdg 3 | 35.2 | 55.8 | 86.4
Ref.inst.rdg 4 | 35.2 | 55.8 | 86.4
Ref.inst.rdg 5 | 35.2 | 55.8 | 86.4
Ref.inst.rdg 6 | 35.2 | 55.8 | 86.4
Ref.inst.rdg 7 | 35.2 | 55.8 | 86.4
Ref.inst.rdg 8 | 35.2 | 55.8 | 86.4
Ref.inst.rdg 9 | 35.2 | 55.8 | 86.4
Ref.inst.rdg 10 | 35.2 | 55.8 | 86.4
MEAN TEMP | 35.20 | 55.80 | 86.40 | 0.00 | 0.00 | 0.00 | 0.00 | 0.00
Standard Deviation Mean | 2.3685e-15 | 0 | 4.737e-15 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
UUT rdg 1 | 35 | 55 | 85
UUT rdg 2 | 35 | 55 | 85
UUT rdg 3 | 35 | 55 | 85
UUT rdg 4 | 35 | 55 | 85
UUT rdg 5 | 35 | 55 | 85
UUT rdg 6 | 35 | 55 | 85
UUT rdg 7 | 35 | 55 | 85
UUT rdg 8 | 35 | 55 | 85
UUT rdg 9 | 35 | 55 | 85
UUT rdg 10 | 35 | 55 | 85
MEAN TEMP | 35.0 | 55.0 | 85.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0
Standard Deviation Mean | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
UNCERTAINTY CALCULATION
DEGREES OF FREEDOM | Infinity |  | Infinity |  | Infinity |  | Infinity |  | Infinity | Infinity |  |  | 9 | 9 | Infinity |  | combined Uncertainty | Eff Degrees of freedom | k-factor @ approx 95% CL (from t-distribution table) | expanded uncertainty | round up value ± | REPORT UNCERTAINTY %r.h.
probability distribution- | Normal |  | Rectangular |  | Normal |  | Rectangular |  | Rectangular | Rectangular |  |  | Normal | Normal | Rectangular
divisor | 2 |  | 1.7321 |  | 2 |  | 1.7321 |  | 1.7321 | 1.7321 |  |  | 1 | 1 | 3.4641 |  |  |  |  | %r.h.
Setpoint | REF Uncertainty %r.h. |  | REF Accuracy %r.h. |  | Indicator Uncertainty %r.h. |  | Indicator Accuracy %r.h. |  | Chamber Stability %r.h. | Chamber Uniformity %r.h. |  |  | Ref Repeatability %r.h. | UUT Repeatability %r.h. | Resolution %r.h.
%r.h.
35.20 | #N/A | #N/A | #N/A | #N/A | 0 | 0 | 0 | 0 | 0.57735 | 0.86603 |  |  | 2.3685e-15 | 0 | 0.1 | 0.028868 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
55.80 | #N/A | #N/A | #N/A | #N/A | 0 | 0 | 0 | 0 | 0.57735 | 0.86603 |  |  | 0 | 0 | 0.1 | 0.028868 | #N/A | #N/A | 2 | #N/A | #N/A | #N/A
86.40 | #N/A | #N/A | #N/A | #N/A | 0 | 0 | 0 | 0 | 0.57735 | 0.86603 |  |  | 4.737e-15 | 0 | 0.1 | 0.028868 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
0.00 | #N/A | #N/A | #N/A | #N/A | 0 | 0 | 0 | 0 | 0.57735 | 0.86603 |  |  | #DIV/0! | #DIV/0! | 0.1 | 0.028868 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
0.00 | #N/A | #N/A | #N/A | #N/A | 0 | 0 | 0 | 0 | 0.57735 | 0.86603 |  |  | #DIV/0! | #DIV/0! | 0.1 | 0.028868 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
0.00 | #N/A | #N/A | #N/A | #N/A | 0 | 0 | 0 | 0 | 0.57735 | 0.86603 |  |  | #DIV/0! | #DIV/0! | 0.1 | 0.028868 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
0.00 | #N/A | #N/A | #N/A | #N/A | 0 | 0 | 0 | 0 | 0.57735 | 0.86603 |  |  | #DIV/0! | #DIV/0! | 0.1 | 0.028868 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
0.00 | #N/A | #N/A | #N/A | #N/A | 0 | 0 | 0 | 0 | 0.57735 | 0.86603 |  |  | #DIV/0! | #DIV/0! | 0.1 | 0.028868 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
Note on G1: DO N'T TYPE THIS ROW
Note on C8: User: select type
Note on C9: User: select type
Note on F98: bs-electrical: /1 IS SENSITIVITY CO EFFICIENT

## Sheet3 (hidden)
[5 empty] | MEASURE MODE |  | REF OUTPUT
|  |  | K TYPE | (K TYPE) | [6 empty] | J TYPE | (J TYPE)
LOGIC | S.NO | TEMPERATURE | UNCERTAINTY | ACCURACY | T/C UC | T/C ERROR |  | LOGIC | S.NO | TEST POINT | UNCERTAINTY | ACCURACY | T/C UC | T/C ERROR
5 | 1 | 1300 | 0.05 | 0.4 | 0.4 | 0.08 |  | 5 | 1 | 1200 | 0.04 | 0.23 | 0.02 | 0.32
5 | 2 | 1000 | 0.05 | 0.4 | 0.4 | 0.08 |  | 5 | 2 | 600 | 0.04 | 0.17 | 0.02 | 0.32
3 | 3 | 600 | 0.05 | 0.26 | 0.4 | 0.08 |  | 2 | 3 | 151 | 0.04 | 0.17 | 0.02 | 0.32
2 | 4 | 200 | 0.05 | 0.26 | 0.4 | 0.08 |  | 1 | 4 | 150 | 0.04 | 0.14 | 0.02 | 0.32
1 | 5 | 50 | 0.05 | 0.16 | 0.4 | 0.08 |  | 1 | 5 | 0 | 0.04 | 0.14 | 0.02 | 0.32
5 | 6 | -200 | 0.05 | 0.33 | 0.4 | 0.08 |  | 5 | 6 | -210 | 0.04 | 0.27 | 0.02 | 0.32
5 | [7 empty] | 5
5 | [7 empty] | 5
5 | [7 empty] | 5
5 | [7 empty] | 5
5 | [7 empty] | 5
5 | [7 empty] | 5
#N/A | [7 empty] | #N/A
5 | [7 empty] | 5
|  |  | T TYPE | (T TYPE) | [6 empty] | R TYPE |  | (R TYPE)
LOGIC | S.NO | TEST POINT | UNCERTAINTY | ACCURACY | T/C UC | T/C ERROR |  | LOGIC | S.NO | TEST POINT | UNCERTAINTY | ACCURACY | T/C UC | T/C ERROR
3 | 1 | 400 | 0.05 | 0.14 | 0.22 | 0.12 |  | 6 | 1 | 1750 | 0.1 | 0.4 | 0.02 | 0.5
3 | 2 | 200 | 0.05 | 0.14 | 0.22 | 0.12 |  | 6 | 2 | 1500 | 0.1 | 0.4 | 0.02 | 0.5
#N/A | 3 | 100 | 0.05 | 0.14 | 0.22 | 0.12 |  | 4 | 3 | 1000 | 0.1 | 0.33 | 0.02 | 0.5
#N/A | 4 | -150 | 0.05 | 0.24 | 0.22 | 0.12 |  | 3 | 4 | 500 | 0.1 | 0.33 | 0.02 | 0.5
#N/A | 5 | -151 | 0.05 | 0.63 | 0.22 | 0.12 |  | 2 | 5 | 250 | 0.1 | 0.35 | 0.02 | 0.5
3 | 6 | -250 | 0.05 | 0.63 | 0.22 | 0.12 |  | 6 | 6 | 0 | 0.1 | 0.57 | 0.02 | 0.5
3 | [7 empty] | 6
3 | [7 empty] | 6
3 | [7 empty] | 6
3 | [7 empty] | 6
3 | [7 empty] | 6
3 | [7 empty] | 6
#N/A | [7 empty] | #N/A
3 | [7 empty] | 6
|  |  | S TYPE | (S TYPE) | [5 empty] | B TYPE | (B TYPE)
LOGIC | S.NO | TEST POINT | UNCERTAINTY | ACCURACY | T/C UC | T/C ERROR |  | LOGIC | S.NO | TEST POINT | UNCERTAINTY | ACCURACY | T/C UC | T/C ERROR
6 | 1 | 1750 | 0.09 | 0.46 | 1.4 | 0.01 |  | 6 | 1 | 1800 | 0.2 | 0.33 | 0.02 | 0.5
6 | 2 | 1500 | 0.09 | 0.46 | 1.4 | 0.01 |  | 6 | 2 | 1500 | 0.2 | 0.3 | 0.02 | 0.5
4 | 3 | 1200 | 0.09 | 0.37 | 1.4 | 0.01 |  | 6 | 3 | 1200 | 0.2 | 0.3 | 0.02 | 0.5
3 | 4 | 600 | 0.09 | 0.36 | 1.4 | 0.01 |  | 4 | 4 | 1000 | 0.2 | 0.3 | 0.02 | 0.5
3 | 5 | 250 | 0.09 | 0.36 | 1.4 | 0.01 |  | 3 | 5 | 800 | 0.2 | 0.44 | 0.02 | 0.5
6 | 6 | 0 | 0.09 | 0.47 | 1.4 | 0.01 |  | 6 | 6 | 600 | 0.2 | 0.44 | 0.02 | 0.5
6 | [7 empty] | 6
6 | [7 empty] | 6
6 | [7 empty] | 6
6 | [7 empty] | 6
6 | [7 empty] | 6
6 | [7 empty] | 6
#N/A | [7 empty] | #N/A
6 | [7 empty] | 6
|  |  | N TYPE | (N TYPE) | [6 empty] | E TYPE | (E TYPE)
LOGIC | S.NO | TEST POINT | UNCERTAINTY | ACCURACY | T/C UC | T/C ERROR |  | LOGIC | S.NO | TEST POINT | UNCERTAINTY | ACCURACY | T/C UC | T/C ERROR
5 | 1 | 1300 | 0.05 | 0.27 | 0.02 | 0.73 |  | 5 | 1 | 1000 | 0.04 | 0.21 | 0.02 | 0.21
5 | 2 | 800 | 0.05 | 0.27 | 0.02 | 0.73 |  | 4 | 2 | 800 | 0.04 | 0.21 | 0.02 | 0.21
2 | 3 | 400 | 0.05 | 0.27 | 0.02 | 0.73 |  | 3 | 3 | 500 | 0.04 | 0.16 | 0.02 | 0.21
1 | 4 | 260 | 0.05 | 0.18 | 0.02 | 0.73 |  | 1 | 4 | 200 | 0.04 | 0.14 | 0.02 | 0.21
1 | 5 | 50 | 0.05 | 0.19 | 0.02 | 0.73 |  | #N/A | 5 | -100 | 0.04 | 0.5 | 0.02 | 0.21
5 | 6 | -200 | 0.05 | 0.4 | 0.02 | 0.73 |  | 4
5 | [7 empty] | 4
5 | [7 empty] | 4
5 | [7 empty] | 4
5 | [7 empty] | 4
5 | [7 empty] | 4
5 | [7 empty] | 4
#N/A | [7 empty] | #N/A
5 | [7 empty] | 4
|  |  | PRT | (PT 100-385)
LOGIC | S.NO | TEST POINT | UNCERTAINTY | ACCURACY | DMM UC
6 | 1 | 800 | 0.2 | 0.23 | 0.1
5 | 2 | 600 | 0.2 | 0.12 | 0.1
2 | 3 | 400 | 0.2 | 0.12 | 0.1
#N/A | 4 | 100 | 0.22 | 0.09 | 0.1
#N/A | 5 | 0 | 0.23 | 0.07 | 0.1
5 | 6 | -100 | 0.25 | 0.05 | 0.1
5
5
5
5
5
5
#N/A
5

## Sheet4 (hidden)
[5 empty] | SOURCE MODE |  | REF INPUT
|  |  | K TYPE | (K TYPE) | [6 empty] | J TYPE | (J TYPE)
LOGIC | S.NO | TEMPERATURE | UNCERTAINTY | ACCURACY | T/C UC | T/C ERROR |  | LOGIC | S.NO | TEST POINT | UNCERTAINTY | ACCURACY | T/C UC | T/C ERROR
5 | 1 | 1300 | 0.07 | 0.4 | 0.4 | 0.08 |  | 5 | 1 | 1200 | 0.06 | 0.23 | 0.02 | 0.12
5 | 2 | 1000 | 0.07 | 0.4 | 0.4 | 0.08 |  | 5 | 2 | 600 | 0.06 | 0.17 | 0.02 | 0.12
3 | 3 | 600 | 0.07 | 0.26 | 0.4 | 0.08 |  | 2 | 3 | 151 | 0.06 | 0.17 | 0.02 | 0.12
2 | 4 | 200 | 0.07 | 0.26 | 0.4 | 0.08 |  | 1 | 4 | 150 | 0.06 | 0.14 | 0.02 | 0.12
1 | 5 | 50 | 0.07 | 0.16 | 0.4 | 0.08 |  | 1 | 5 | 0 | 0.06 | 0.14 | 0.02 | 0.12
5 | 6 | -200 | 0.07 | 0.33 | 0.4 | 0.08 |  | 5 | 6 | -210 | 0.06 | 0.27 | 0.02 | 0.12
5 | [7 empty] | 5
5 | [7 empty] | 5
5 | [7 empty] | 5
5 | [7 empty] | 5
5 | [7 empty] | 5
5 | [7 empty] | 5
#N/A | [7 empty] | #N/A
5 | [7 empty] | 5
|  |  | T TYPE | (T TYPE) | [6 empty] | R TYPE |  | (R TYPE)
LOGIC | S.NO | TEST POINT | UNCERTAINTY | ACCURACY | T/C UC | T/C ERROR |  | LOGIC | S.NO | TEST POINT | UNCERTAINTY | ACCURACY | T/C UC | T/C ERROR
3 | 1 | 400 | 0.05 | 0.14 | 0.22 | 0.12 |  | 6 | 1 | 1750 | 0.12 | 0.4 | 0.02 | 0.5
3 | 2 | 200 | 0.05 | 0.14 | 0.22 | 0.12 |  | 6 | 2 | 1500 | 0.12 | 0.4 | 0.02 | 0.5
#N/A | 3 | 100 | 0.05 | 0.14 | 0.22 | 0.12 |  | 4 | 3 | 1000 | 0.12 | 0.33 | 0.02 | 0.5
#N/A | 4 | -150 | 0.05 | 0.24 | 0.22 | 0.12 |  | 3 | 4 | 500 | 0.12 | 0.33 | 0.02 | 0.5
#N/A | 5 | -151 | 0.05 | 0.63 | 0.22 | 0.12 |  | 2 | 5 | 250 | 0.12 | 0.35 | 0.02 | 0.5
3 | 6 | -250 | 0.05 | 0.63 | 0.22 | 0.12 |  | 6 | 6 | 0 | 0.12 | 0.57 | 0.02 | 0.5
3 | [7 empty] | 6
3 | [7 empty] | 6
3 | [7 empty] | 6
3 | [7 empty] | 6
3 | [7 empty] | 6
3 | [7 empty] | 6
#N/A | [7 empty] | #N/A
3 | [7 empty] | 6
|  |  | S TYPE | (S TYPE) | [5 empty] | B TYPE | (B TYPE)
LOGIC | S.NO | TEST POINT | UNCERTAINTY | ACCURACY | T/C UC | T/C ERROR |  | LOGIC | S.NO | TEST POINT | UNCERTAINTY | ACCURACY | T/C UC | T/C ERROR
6 | 1 | 1750 | 0.12 | 0.46 | 1.4 | 0.01 |  | 6 | 1 | 1800 | 0.2 | 0.33 | 0.02 | 0.5
6 | 2 | 1500 | 0.12 | 0.46 | 1.4 | 0.01 |  | 6 | 2 | 1500 | 0.2 | 0.3 | 0.02 | 0.5
4 | 3 | 1200 | 0.12 | 0.37 | 1.4 | 0.01 |  | 6 | 3 | 1200 | 0.2 | 0.3 | 0.02 | 0.5
3 | 4 | 600 | 0.12 | 0.36 | 1.4 | 0.01 |  | 4 | 4 | 1000 | 0.2 | 0.3 | 0.02 | 0.5
3 | 5 | 250 | 0.12 | 0.36 | 1.4 | 0.01 |  | 3 | 5 | 800 | 0.2 | 0.44 | 0.02 | 0.5
6 | 6 | 0 | 0.12 | 0.47 | 1.4 | 0.01 |  | 6 | 6 | 600 | 0.2 | 0.44 | 0.02 | 0.5
6 | [7 empty] | 6
6 | [7 empty] | 6
6 | [7 empty] | 6
6 | [7 empty] | 6
6 | [7 empty] | 6
6 | [7 empty] | 6
#N/A | [7 empty] | #N/A
6 | [7 empty] | 6
|  |  | N TYPE | (N TYPE) | [6 empty] | E TYPE | (E TYPE)
LOGIC | S.NO | TEST POINT | UNCERTAINTY | ACCURACY | T/C UC | T/C ERROR |  | LOGIC | S.NO | TEST POINT | UNCERTAINTY | ACCURACY | T/C UC | T/C ERROR
5 | 1 | 1300 | 0.06 | 0.27 | 0.02 | 0.73 |  | 5 | 1 | 1000 | 0.05 | 0.21 | 0.02 | 0.21
5 | 2 | 800 | 0.06 | 0.27 | 0.02 | 0.73 |  | 4 | 2 | 800 | 0.05 | 0.21 | 0.02 | 0.21
2 | 3 | 400 | 0.06 | 0.27 | 0.02 | 0.73 |  | 3 | 3 | 500 | 0.05 | 0.16 | 0.02 | 0.21
1 | 4 | 260 | 0.06 | 0.18 | 0.02 | 0.73 |  | 1 | 4 | 200 | 0.05 | 0.14 | 0.02 | 0.21
1 | 5 | 50 | 0.06 | 0.19 | 0.02 | 0.73 |  | #N/A | 5 | -100 | 0.05 | 0.5 | 0.02 | 0.21
5 | 6 | -200 | 0.06 | 0.4 | 0.02 | 0.73 |  | 4
5 | [7 empty] | 4
5 | [7 empty] | 4
5 | [7 empty] | 4
5 | [7 empty] | 4
5 | [7 empty] | 4
5 | [7 empty] | 4
#N/A | [7 empty] | #N/A
5 | [7 empty] | 4
|  |  | PRT | (PT 100-385)
LOGIC | S.NO | TEST POINT | UNCERTAINTY | ACCURACY
6 | 1 | 800 | 0.2 | 0.08
5 | 2 | 600 | 0.2 | 0.08
2 | 3 | 400 | 0.2 | 0.07
#N/A | 4 | 100 | 0.2 | 0.07
#N/A | 5 | 0 | 0.2 | 0.07
5 | 6 | -100 | 0.2 | 0.01
5
5
5
5
5
5
#N/A
5

## REFERANCE
password bstech
S/N | S/N | Tag no. | Description | Serial no. | Cal date | Due Date | Section | Date
1 | 1 | BS 1320 | Multifunction Calibrator | 107247G5 | 2012-08-03 | 2014-08-03 | BSE 01 or BST 05 | FAIL | 2012-08-03 | 2014-08-03 | BS 1320 | 0
2 | 2 | BS 1308 | Precision DMM | US36085948 | 2012-11-07 | 2013-11-07 | BSE 02 or BST 05 | FAIL | 2012-11-07 | 2013-11-07 | BS 1308 | 0
3 | 3 | BS 1326 | Precision DMM | US36075261 | 2013-10-03 | 2014-10-03 | BSE 02 or BST 05 | PASS | 2013-10-03 | 2014-10-03 | BS 1326 | 0
4 | 4 | BS 1306 | D P Calibrator | 7849010 | 2013-04-15 | 2014-04-15 | BSE 01 & BST 05 | FAIL | 2013-04-15 | 2014-04-15 | BS 1306 | 0
5 | 5 | BS 1307 | High Voltage probe | Z4/6625-01-193-1006 | 2012-03-28 | 2013-03-28 | BSE 04 | FAIL | 2012-03-28 | 2013-03-28 | BS 1307 | 0
6 | 6 | BS 1304 | Decade Resistance Box | 064066 | 2011-10-31 | 2013-10-31 | BSE 03 | FAIL | 2011-10-31 | 2013-10-31 | BS 1304 | 0
7 | 7 | BS 1310 | Current Shunt | - | 2012-04-09 | 2013-04-09 | BSE 05 | FAIL | 2012-04-09 | 2013-04-09 | BS 1310 | 0
8 | 8 | BS 1364 | Low/High Resistance Box | R0001 | 2012-11-14 | 2013-11-14 | BSE 03 | FAIL | 2012-11-14 | 2013-11-14 | BS 1364 | 0
9 | 9 | BS 1394 | Multifunction Calibrator | 9760013 | 2013-08-02 | 2014-08-02 | BST 05 | FAIL | 2013-08-02 | 2014-08-02 | BS 1394 | 0
10 | 10 | BS 1409 | High Precision DMM | MY45043681 | 2011-09-30 | 2013-09-30 | BSE 02 | FAIL | 2011-09-30 | 2013-09-30 | BS 1409 | 0
11 | 11 | BS 1365 | High Voltage Resistance Box | R0003 | 2012-10-23 | 2013-09-03 | BSE 03 | FAIL | 2012-10-23 | 2013-09-03 | BS 1365 | 0
12 | 12 | BS 1309 | Current Shunt (30A) | 3812-04-0344-01 | 2012-01-25 | 2013-01-25 | BSE 05 | FAIL | 2012-01-25 | 2013-01-25 | BS 1309 | 0
13 | 13 | BS 1330 | Standard Resistance Box | C82189 | 2011-01-19 | 2012-01-19 | BSE 03 | FAIL | 2011-01-19 | 2012-01-19 | BS 1330 | 0
14 | 14 | BS 1331 | Clamp Meter | 97AM2089 | 2012-01-25 | 2013-01-25 | BSE 02 | FAIL | 2012-01-25 | 2013-01-25 | BS 1331 | 0
15 | 15 | BS 1336 | Hi-Pot Tester | 2146 | 2011-11-17 | 2012-11-17 | BSE 04 | FAIL | 2011-11-17 | 2012-11-17 | BS 1336 | 0
16 | 16 | BS 1327 | Power Supply | 2936K 01148 | 2012-01-25 | 2013-01-25 | BSE 01 | FAIL | 2012-01-25 | 2013-01-25 | BS 1327 | 0
17 | 17 | BS 1332 | DC Power Supply | 0070/07 | 2012-01-18 | 2013-01-18 | BSE 01 | FAIL | 2012-01-18 | 2013-01-18 | BS 1332 | 0
18 | 18 | BS 1396 | Decade Resistance Box | 66VX0199 | 2011-10-24 | 2012-10-24 | BSE 03 | FAIL | 2011-10-24 | 2012-10-24 | BS 1396 | 0
19 | 19 | BS 1397 | Digital Multimeter | 76255611 | 2012-09-01 | 2013-09-01 | BSE 02 | FAIL | 2012-09-01 | 2013-09-01 | BS 1397 | 0
20 | 20 | BS 1398 | GHz Multifunction Counter | 98110041 | 2011-05-28 | 2012-05-28 | BSE 02 | FAIL | 2011-05-28 | 2012-05-28 | BS 1398 | 0
21 | 21 | BS 1399 | AC/DC Clamp Meter | W8049010 | 2011-11-15 | 2012-11-15 | BSE 02 | FAIL | 2011-11-15 | 2012-11-15 | BS 1399 | 0
22 | 22 | BS 1402 | MULTIFUNCTION CALIBRATOR | 6325016 | 2011-10-27 | 2012-10-27 | BSE 01 | FAIL | 2011-10-27 | 2012-10-27 | BS 1402 | 0
23 | 23 | BS 1446 | Digital Timer | 2355059 | 2011-10-25 | 2012-10-25 | BSE 07 | FAIL | 2011-10-25 | 2012-10-25 | BS 1446 | 0
24 | 1 | BS 1317 | Pneumatic Pressure Calibrator | 8893062 | 2012-07-30 | 2013-07-30 | BSM 02 | FAIL | 2012-07-30 | 2013-07-30 | BS 1317 | 0
25 | 2 | BS 1319 | Low Pressure/Vacc. calibrator | 8882036 | 2012-07-30 | 2013-07-30 | BSM 03 | FAIL | 2012-07-30 | 2013-07-30 | BS 1319 | 0
26 | 3 | BS 1316 | Absolute Transducer | 5134 | 2012-04-16 | 2013-04-16 | BSM 06 | FAIL | 2012-04-16 | 2013-04-16 | BS 1316 | 0
27 | 4 | BS 1312 | Dead Weight Tester | 580/28456 | 2012-05-07 | 2015-05-07 | BSM 01 | PASS | 2012-05-07 | 2015-05-07 | BS 1312 | 0
28 | 5 | BS 1312 | Weights | A8216 | 2012-05-02 | 2013-05-02 | BSM 01 | FAIL | 2012-05-02 | 2013-05-02 | BS 1312 | 0
29 | 6 | BS 1360 | High Vacuum System | 1730, 350 | 2011-11-22 | 2013-11-22 | BSM 05 | FAIL | 2011-11-22 | 2013-11-22 | BS 1360 | 0
30 | 7 | BS 1314 | Low Pressure Calibrator | - | 2012-07-23 | 2013-07-23 | BSM 04 | FAIL | 2012-07-23 | 2013-07-23 | BS 1314 | 0
31 | 8 | BS 1355 | Pressure Transducer | 251 | 2012-09-05 | 2013-09-05 | BSM 01 | FAIL | 2012-09-05 | 2013-09-05 | BS 1355 | 0
32 | 9 | BS 1356 | Pressure Transducer | 3885 | 2012-08-01 | 2013-08-01 | BSM 02 | FAIL | 2012-08-01 | 2013-08-01 | BS 1356 | 0
33 | 10 | BS 1351 | Compound Transducer | 3908 | 2012-09-05 | 2013-09-05 | BSM 03 | FAIL | 2012-09-05 | 2013-09-05 | BS 1351 | 0
34 | 11 | BS 1390 | Pressure Calibrator | 60306137 | 2012-09-05 | 2013-09-05 | BSM 02 | FAIL | 2012-09-05 | 2013-09-05 | BS 1390 | 0
35 | 12 | DWT USE | Thermocouple type-k | - | 2013-03-28 | 2014-03-28 | BST 05 | FAIL | 2013-03-28 | 2014-03-28 | DWT USE | 0
36 | 13 | BS 1442 | Pressure Calibrator | 2360-840650 | 2012-09-05 | 2013-09-05 | BSM 03 | FAIL | 2012-09-05 | 2013-09-05 | BS 1442 | 0
37 | 14 | BS 1443 | Pressure Calibrator | 869181 | 2012-09-05 | 2013-09-05 | BSM 01 | FAIL | 2012-09-05 | 2013-09-05 | BS 1443 | 0
38 | 15 | BS 1444 | Pressure Calibrator | 869634 | 2012-10-29 | 2013-10-29 | BSM 02 | FAIL | 2012-10-29 | 2013-10-29 | BS 1444 | 0
39 | 16 | BS 1339 | Hydraulic valve test bench | 6863 | 2012-09-05 | 2013-09-05 | BSM 09 | FAIL | 2012-09-05 | 2013-09-05 | BS 1339 | 0
40 | 17 | BS 1323 | HANDHELD tachometer | 4310083 | 2012-03-09 | 2013-03-09 | BSM 08 | FAIL | 2012-03-09 | 2013-03-09 | BS 1323 | 0
41 | 18 | BS 1460 | High Pressure Calibrator | 2028005885 | 2012-10-09 | 2013-10-09 | BSM 01 | FAIL | 2012-10-09 | 2013-10-09 | BS 1460 | 0
42 | 1 | BS 1328 | SPRT | 088 | 2013-11-25 | 2015-11-25 | BST 01-H/L | PASS | 2013-11-25 | 2015-11-25 | BS 1328 | 0
43 | 2 | BS 1313 | Standard Thermocouple | 25622/1 | 2013-05-27 | 2014-05-27 | BST 02 | FAIL | 2013-05-27 | 2014-05-27 | BS 1313 | 0
44 | 3 | BS 1340 | PRT SENSOR(PT 100)-4 WIRE | EP 53022 | 2013-06-20 | 2014-06-20 | BST 01-H/L | FAIL | 2013-06-20 | 2014-06-20 | BS 1340 | 0
45 | 4 | BS 1321 | Digital ThermoHygrometer | 37294 029/60587499 | 2011-08-08 | 2012-08-08 | BST 07 | FAIL | 2011-08-08 | 2012-08-08 | BS 1321 | 0
46 | 5 | BS 1393 | Digital ThermoHygrometer | 37294 029/60587498 | 2013-08-22 | 2014-08-22 | BST 07 | FAIL | 2013-08-22 | 2014-08-22 | BS 1393 | 0
47 | 6 | BS 1322 | Infrared Thermometer | 212220-0301-0008 | 2011-07-18 | 2013-07-18 | BST 03 | FAIL | 2011-07-18 | 2013-07-18 | BS 1322 | 0
48 | 7 | BS 1311 | Standard Thermocouple with indicator | 25622/2 / A43620 | 2013-08-19 | 2014-08-19 | BST 02 | FAIL | 2013-08-19 | 2014-08-19 | BS 1311 | 0
49 | 8 | BS 1400 | Digital Thermometer | A43620 | 2011-11-22 | 2013-11-22 | BST 05 | FAIL | 2011-11-22 | 2013-11-22 | BS 1400 | 0
50 | 9 | BS 1358 | Type-K Thermocouple (H) | - | 2013-06-20 | 2014-06-20 | BST 06 | FAIL | 2013-06-20 | 2014-06-20 | BS 1358 | 0
51 | 10 | BS 1359 | Type-K Thermocouple (L) | - | 2013-03-28 | 2014-03-28 | BST 06 | FAIL | 2013-03-28 | 2014-03-28 | BS 1359 | 0
52 | 11 | BS 1362 | Data Logger | AY438/759 | 2013-03-28 | 2014-03-28 | BST 06 | FAIL | 2013-03-28 | 2014-03-28 | BS 1362 | 0
53 | 12 | BS 1363 | Reference thermocouple | - | 2013-10-24 | 2014-10-24 | BST 05 | PASS | 2013-10-24 | 2014-10-24 | BS 1363 | 0
54 | 13 | BS 1343 | Type-k temperature probe | K1 | 2012-12-03 | 2013-12-03 | BST 02 | FAIL | 2012-12-03 | 2013-12-03 | BS 1343 | 0
55 | 14 | BS 1324 | Thermo hygrograph | 40260205 | 2012-12-07 | 2013-12-07 | BST 07 | FAIL | 2012-12-07 | 2013-12-07 | BS 1324 | 0
56 | 15 | BS 8028 | Thermo hygrograph | 0229722-24 | 2012-07-25 | 2013-07-25 | BST 07 | FAIL | 2012-07-25 | 2013-07-25 | BS 8028 | 0
57 | 16 | BS 1439 | Thermo hygrograph | 0229722-25 | 2012-09-10 | 2013-09-10 | BST 07 | FAIL | 2012-09-10 | 2013-09-10 | BS 1439 | 0
58 | 17 | BS 1346 | Black Body Calibrator | 6010103 | 2012-05-11 | 2013-05-11 | BST 03 | FAIL | 2012-05-11 | 2013-05-11 | BS 1346 | 0
59 | 18 | BS 1391 | Dry Block Calibrator | 241016 | 2012-06-04 | 2013-06-04 | BST 01-H & BST 02 | FAIL | 2012-06-04 | 2013-06-04 | BS 1391 | 0
60 | 19 | BS 1342 | Mid. Range Temp. Calibrator | FB0510 10596 | 2012-06-01 | 2013-06-01 | BST 01-H & BST 02 | FAIL | 2012-06-01 | 2013-06-01 | BS 1342 | 0
61 | 20 | BS 1344 | Liquid Temp. Bath | IRI 0510 10597 | 2013-04-06 | 2014-04-06 | BST 01-H/L | FAIL | 2013-04-06 | 2014-04-06 | BS 1344 | 0
62 | 21 | BS 1335 | Temp. & Humidity Chamber | 6458EC | 2013-02-14 | 2014-02-14 | BST 07 | FAIL | 2013-02-14 | 2014-02-14 | BS 1335 | 0
63 | 22 | BS 1347 | Tube Furnace Calibrator | 1260050890-190 | 2013-02-07 | 2014-02-07 | BST 02 | FAIL | 2013-02-07 | 2014-02-07 | BS 1347 | 0
64 | 23 | BS 1338 | Hot Plate Calibrator | 1501 | 2012-05-31 | 2013-05-31 | BST 08 | FAIL | 2012-05-31 | 2013-05-31 | BS 1338 | 0
65 | 24 | BS 1407 | Hygrogen Calibrator | 51967001 | 2012-05-04 | 2013-05-04 | BST 07 | FAIL | 2012-05-04 | 2013-05-04 | BS 1407 | 0
66 | 25 | BS 1383 | Black Body Calibrator | 05558 | 2012-06-04 | 2013-06-04 | BST 03 | FAIL | 2012-06-04 | 2013-06-04 | BS 1383 | 0
67 | 26 | BS 1463 | MULTIFUNCTIONAL TEMP. CALIBRATOR | 31611020010 | 2013-04-05 | 2014-04-05 | BST 05 | FAIL | 2013-04-05 | 2014-04-05 | BS 1463 | 0
67 | 27 | BS 1467 | FADEN THERMOMETER | 3461 | 2013-07-08 | 2014-07-08 | BST 01 L/H | FAIL | 2013-07-08 | 2014-07-08 | BS 1467 | 0
68 | 28 | BS 1318 | ICE POINT | - | - | - | BST 04 | PASS |  |  | BS 1318 | 0
69 | 1 | BS 1302 | Acoustical Calibrator | 030704010 | 2012-01-27 | 2013-01-27 | BSMS 06-2008 | FAIL | 2012-01-27 | 2013-01-27 | BS 1302 | 0
70 | 2 | BS 1449-1 | pH Standard cal. SOLUTION | - | 2011-02-18 | 2012-02-18 | BSMS 04-2008 | FAIL | 2011-02-18 | 2012-02-18 | BS 1449-1 | 0
71 | 3 | BS 1449-2 | CONDU. standard CAL. SOLUTION | - | 2011-09-16 | 2012-09-16 | BSMS 04-2008 | FAIL | 2011-09-16 | 2012-09-16 | BS 1449-2 | 0
72 | 4 | BS 1449-3 | TDS standard CAL. SOLUTION | 2253 | 2011-03-09 | 2012-03-09 | BSMS 04-2008 | FAIL | 2011-03-09 | 2012-03-09 | BS 1449-3 | 0
73 | 5 | BS 1449-4 | ORP standard CAL. SOLUTION | 1256 | 2010-12-03 | 2011-12-03 | BSMS 04-2008 | FAIL | 2010-12-03 | 2011-12-03 | BS 1449-4 | 0
74 | 6 | BS 1449-5 | SALT standard CAL. SOLUTION | 0096 | 2011-09-11 | 2012-09-11 | BSMS 04-2008 | FAIL | 2011-09-11 | 2012-09-11 | BS 1449-5 | 0
75 | 7 | BS 1369 | Load Cell with indicator | 606375 | 2011-07-19 | 2012-07-19 | BSMS 01-2008 | FAIL | 2011-07-19 | 2012-07-19 | BS 1369 | 0
76 | 8 | BS 1376 | White light meter | Q290048 | 2012-02-10 | 2013-02-10 | BSMS 02-2008 | FAIL | 2012-02-10 | 2013-02-10 | BS 1376 | 0
77 | 9 | BS 1377 | Ultra violet meter/ultra violet lamp | C66505 | 2010-12-24 | 2011-12-24 | BSMS 02-2008 | FAIL | 2010-12-24 | 2011-12-24 | BS 1377 | 0
78 | 10 | BS 1403 | Sound Level Meter | 1431 | 2012-01-12 | 2013-01-12 |  | FAIL | 2012-01-12 | 2013-01-12 | BS 1403 | 0
79 | 11 | BS 1371 | Flow calibrator | G79743-8 | 2011-07-18 | 2012-07-18 |  | FAIL | 2011-07-18 | 2012-07-18 | BS 1371 | 0
80 | 12 | BS 1368 | Thermo-HOT WIRE ANEMOMETER | Q412901 | 2011-02-10 | 2012-02-10 |  | FAIL | 2011-02-10 | 2012-02-10 | BS 1368 | 0
81 | 13 | BS 1462 | ANEMOMETER | 11043088 | 2011-08-20 | 2012-08-20 |  | FAIL | 2011-08-20 | 2012-08-20 | BS 1462 | 0
82 | 1 | BS 8001 | GLASS SCALE | 12048 | 2010-08-13 | 2013-08-13 |  | FAIL | 2010-08-13 | 2013-08-13 | BS 8001 | 0
83 | 2 | BS 8002 | PEAK GLASS SCALE | - | 2010-08-13 | 2013-08-13 |  | FAIL | 2010-08-13 | 2013-08-13 | BS 8002 | 0
84 | 3 | BS 8003 | PRECISION BALL | - | 2010-12-27 | 2011-12-27 |  | FAIL | 2010-12-27 | 2011-12-27 | BS 8003 | 0
85 | 4 | BS 8005 | DIGIMATIC MICROMETER | 85088848 | 2011-09-27 | 2012-09-27 |  | FAIL | 2011-09-27 | 2012-09-27 | BS 8005 | 0
86 | 5 | BS 8006 | THREAD WIRES | 1273 | 2011-01-04 | 2012-01-04 |  | FAIL | 2011-01-04 | 2012-01-04 | BS 8006 | 0
87 | 6 | BS 8009 | GAUGE BLOCK | 080199 | 2010-11-11 | 2011-11-11 |  | FAIL | 2010-11-11 | 2011-11-11 | BS 8009 | 0
88 | 7 | BS 8010 | GAUGE BLOCK | 01588 | 2010-11-26 | 2011-11-26 |  | FAIL | 2010-11-26 | 2011-11-26 | BS 8010 | 0
89 | 8 | BS 8015 | SURFACE ROUGHNESS SPECIMEN | - | 2011-11-01 | 2012-11-01 |  | FAIL | 2011-11-01 | 2012-11-01 | BS 8015 | 0
90 | 9 | BS 8016 | GAUGE BLOCK | 765458 | 2010-12-13 | 2011-12-13 |  | FAIL | 2010-12-13 | 2011-12-13 | BS 8016 | 0
91 | 10 | BS 8019 | OPTICAL PARALLELS | - | 2010-08-13 | 2015-08-13 |  | PASS | 2010-08-13 | 2015-08-13 | BS 8019 | 0
92 | 11 | BS 8021 | OPTICAL FLAT | 4A4 | 2008-06-26 | 2013-06-26 |  | FAIL | 2008-06-26 | 2013-06-26 | BS 8021 | 0
93 | 12 | BS 8022 | SUPRAMESS INDICATOR | 845754 | 2011-08-10 | 2012-08-10 |  | FAIL | 2011-08-10 | 2012-08-10 | BS 8022 | 0
94 | 13 | BS 8023 | DIAL TEST INDICATOR | T22875 | 2011-08-10 | 2012-08-10 |  | FAIL | 2011-08-10 | 2012-08-10 | BS 8023 | 0
95 | 14 | BS 8024 | DIAL TEST INDICATOR | GNN709 | 2011-08-19 | 2012-08-19 |  | FAIL | 2011-08-19 | 2012-08-19 | BS 8024 | 0
96 | 15 | BS 8026 | SURFACE TABLE | - | 2010-08-17 | 2013-08-17 |  | FAIL | 2010-08-17 | 2013-08-17 | BS 8026 | 0
97 | 16 | BS 8030 | RING GAUGE | 6407 | 2011-08-17 | 2012-08-17 |  | FAIL | 2011-08-17 | 2012-08-17 | BS 8030 | 0
98 | 17 | BS 8034 | ANGLE BLOCK GAUGE | - | 2010-11-18 | 2011-11-18 |  | FAIL | 2010-11-18 | 2011-11-18 | BS 8034 | 0
99 | 18 | BS 8034 | ANGLE BLOCK GAUGE | - | 2010-11-18 | 2011-11-18 |  | FAIL | 2010-11-18 | 2011-11-18 | BS 8034 | 0
100 | 19 | BS 8035 | SMARTSCOPE | S12061632 | 2011-02-28 | 2012-02-28 |  | FAIL | 2011-02-28 | 2012-02-28 | BS 8035 | 0
101 | 20 | BS 8036 | CONCENTRICITY GAUGE | - | 2010-09-30 | 2013-09-30 |  | FAIL | 2010-09-30 | 2013-09-30 | BS 8036 | 0
102 | 21 | BS 8039 | DIGIMATIC CALIPER | 040494 | 2011-09-27 | 2012-09-27 |  | FAIL | 2011-09-27 | 2012-09-27 | BS 8039 | 0
103 | 22 | BS 8048 | STEEL RULE | - | 2011-09-26 | 2012-09-26 |  | FAIL | 2011-09-26 | 2012-09-26 | BS 8048 | 0
104 | 23 | BS 8049 | COATING THICKNESS SPECIMEN | - | 2010-09-24 | 2012-09-24 |  | FAIL | 2010-09-24 | 2012-09-24 | BS 8049 | 0
105 | 24 | BS 8050 | MICROMETER HEAD | 8018582 | 2010-12-30 | 2011-12-30 |  | FAIL | 2010-12-30 | 2011-12-30 | BS 8050 | 0
106 | 25 | BS 8008 | ULM | - | - | - |  | #VALUE! | - | - | BS 8008 | 0
107 | 26 | BS 8011 | LENGTH BARS | - | - | - |  | #VALUE! | - | - | BS 8011 | 0
108 | 27 | BS 8027 | LEVEL GAUGE CALIBRATOR | - | - | - |  | #VALUE! | - | - | BS 8027 | 0
109 | 28 | BS 8031 | PLAIN RING GAUGES | - | - | - |  | #VALUE! | - | - | BS 8031 | 0
110 | 29 | BS 8037 | STEEL MICROMETER SET | - | - | - |  | #VALUE! | - | - | BS 8037 | 0
111 | 30 | BS 8043 | MIC PIN GAUGE | - | - | - |  | #VALUE! | - | - | BS 8043 | 0
112 | 31 | BS 8052 | GAUGE BLOCK | 110006 | - | - |  | #VALUE! | - | - | BS 8052 | 0
113 | 1 | BS 1423 | ULTRA STANDARD WEIGHT | 61260 | 2010-11-12 | 2011-11-12 |  | FAIL | 2010-11-12 | 2011-11-12 | BS 1423 | BS 1423
114 | 2 | BS 1428 | STAINLESS STEEL WEIGHTS | G0812354 | 2011-11-22 | 2012-11-22 |  | FAIL | 2011-11-22 | 2012-11-22 | BS 1428 | BS 1428
115 | 3 | BS 1425 | PRECISION BALANCE | 087790075 | 2011-10-04 | 2012-10-04 |  | FAIL | 2011-10-04 | 2012-10-04 | BS 1425 | BS 1425
116 | 4 | BS 1426 | PRECISION BALANCE | W085073 | 2011-10-04 | 2012-10-04 |  | FAIL | 2011-10-04 | 2012-10-04 | BS 1426 | BS 1426
117 | 5 | BS 1427 | PRECISION BALANCE | W085007 | 2011-10-04 | 2012-10-04 |  | FAIL | 2011-10-04 | 2012-10-04 | BS 1427 | BS 1427
118 | 6 | BS 1440 | PRECISION BALANCE | 8026501101 | 2011-10-04 | 2012-10-04 |  | FAIL | 2011-10-04 | 2012-10-04 | BS 1440 | BS 1440
119 | 7 | BS 1372 | CDI TORQUE TOOL TESTER | 037 | 2011-05-10 | 2012-05-10 |  | FAIL | 2011-05-10 | 2012-05-10 | BS 1372 | BS 1372
120 | 8 | BS 1429 | DUROMETER TESTER | 105026 | 2011-07-07 | 2012-07-07 |  | FAIL | 2011-07-07 | 2012-07-07 | BS 1429 | BS 1429
121 | 9 | BS 1382 | STANDARD MASS SET | - | 2011-08-08 | 2012-08-08 |  | FAIL | 2011-08-08 | 2012-08-08 | BS 1382 | BS 1382
122 | 10 | BS 1436 | STANDARD MASS SET | - | 2011-06-18 | 2012-06-18 |  | FAIL | 2011-06-18 | 2012-06-18 | BS 1436 | BS 1436
| 125 | BS 1318 | STAINLESS STEEL WEIGHT | - | 2009-10-31 | 2011-10-31 | BSM 07
| 126 | BS 1329 | STANDARD MASS SET | 11118462 | 2009-05-10 | 2011-05-10 | BSM 07
| 127 | BS 1382 | STANDARD MASS SET | - | 2009-07-29 | 2011-07-29 | BSM 07
| 128 | BS 1436 | STANDARD MASS SET | - | 2009-05-21 | 2011-05-21 | BSM 07
| 129 | BS 1436-1 | STANDARD MASS SET | - | 2009-05-21 | 2011-05-21 | BSM 07
| 130 | BS 1425 | PRECISION BALANCE | 087790075 | 2010-09-13 | 2011-09-13 | BSM 07
| 131 | BS 1426 | PRECISION BALANCE | w085073 | 2010-09-13 | 2011-09-13 | BSM 07
| 132 | BS 1427 | PRECISION BALANCE | W085007 | 2010-09-13 | 2011-09-13 | BSM 07
| 133 | BS 1440 | PRECISION BALANCE | 8026501101 | 2010-09-13 | 2011-09-13 | BSM 07
| 134 | BS 1372 | CDI TORQUE TOOL TESTER | - | 2010-05-12 | 2011-05-12 | BSM 15
| 135 | BS 1417 | CALIBRATION TORQUE ARM | - |  |  | BSM 11
| 136 | BS 1372 | CALIBRATION TORQUE ARM | - |  |  | BSM 11
| 137 | BS 1422 | MEASURING CYLINDER | -
| 138 | BS 1419 | MEASURING CYLINDER | -
| 139 | BS 1429 | DUROMETER DEADWEIGHTS | 105026 |  |  | BSM 14
| 140 | BS 3005 | DIAL TEST INDICATOR | 3F26387 |  |  | BSD-07-2008
141
Formulas:
Cal date = '[5]Master list-Mechanical-new'!$J13
Due Date = '[5]Master list-Mechanical-new'!$K13
Note on F130: USER06: KGF
Note on F131: USER06: LBF